## Benchsheet
BOD Benchsheet [version 12/4/23] |  |  |  | Facility: | [5 empty] | Method Number: | SM 5210B - 2001
Sampling Information
Sample ID | Collect Date | Collect Time | Sample pH (SU) | pH if adjusted (SU) | Sample Temp Prior to Dilution (°C) [17-23°C] |  | Res. Cl Check (+/-)* | Res. Cl Treated (y/n) |  | Read In Calibration |  | Read Out Calibration
[10 empty] | Analyst |  | Analyst
[10 empty] | Barom. Pressure (mm Hg) |  | Barom. Pressure (mm Hg)
[10 empty] | Room Temp. (°C) |  | Room Temp. (°C)
[10 empty] | DO Saturation Point (mg/L) |  | DO Saturation Point (mg/L)
[10 empty] | Meter Calibration DO (mg/L) |  | Meter Calibration DO (mg/L)
[10 empty] | Date Bottles In |  | Date Bottles Out
[10 empty] | Time Bottles In |  | Time Bottles Out
[6 empty] | * | "+" = checked and res. chlorine is present
Traceability Information | [6 empty] | "-" = checked and res. chlorine is not present
Seed Source Date/Lot:
GGA: | [9 empty] | Quality Control
PHOSPHATE BUFFER: |  |  |  |  | Code Definitions |  |  |  |  | Blanks ≤ 0.24 mg/L |  | Residual DO ≥ 1.0 mg/L
MAGNESIUM SULFATE: |  |  |  |  | N = Extra Nutrients added to sample |  |  |  |  | GGA = 167.5 - 228.5 |  | DO Depletion ≥ 2.0 mg/L
CALCIUM CHLORIDE: |  |  |  |  | I = Inhibitor added to sample |  |  |  |  | pH = 6.0 - 8.5; if not, adjust to 6.5 - 7.5 |  | Sample and Room Temp 20 ± 3 °C
FERRIC CHLORIDE: |  |  |  |  | P = Sample pre-diluted per SOP (≤ 3 mL) |  |  |  |  | Residual Chlorine < 0.1 mg/L |  | Incubator Temp 20 ± 1 °C
SODIUM SULFITE:
Sample ID (must be unique) | Bottle ID | Sample Volume Used (mL) | Code | Pre-dilution Factor (only use if different than 1) | Seed Volume Added (mL) | Initial DO (mg/L) | Final DO (mg/L) | DO Depletion (mg/L) | Sample Seed CF (mg/L) | Seed Adjusted DO Depletion ((mg/L)/mL) | Dilution Factor = 300/Sample Volume Used (mL) | BOD (mg/L) | Average BOD for dilutions that passed (mg/L) | BOD (mg/L) if DO Depl = 2 | BOD (mg/L) if Final DO = 1 | All Deplete <2? | All Final <1? | RL if all Deplete <2 | RL if all Final DO <1 | Max volume/ Corr. BOD if all <2 | Min volume/ Corr. BOD if all <1
Blank |  | 300
Seed Control 1 |  | 300 | [11 empty] | #VALUE! | all text | calc. with passing depletion results | calc. with passing Final DO results
Seed Control 2 |  | 300 | [11 empty] | #VALUE!
Seed Control 3 | [13 empty] | #VALUE!
GGA |  | 6
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
[16 empty] | calc. with passing depletion results | calc. with passing Final DO results | #N/A | #N/A | 0 | 0
[20 empty] | #N/A | #N/A
COMMENTS:
Be sure to clear all previous data before adding new data.
Sample pH cells will turn red or blue if the pH needs to be adjusted.
Temperature cells will turn red or blue if the temperature is not within the required range of 17 - 23°C.
If the sample was chlorinated, the sample must be checked for residual chlorine with strips or another method that can detect down to 0.1 mg/L.
Meter Calibration DO cells will turn red if the value is more than 0.2 mg/L different from the DO saturation point. Recalibrate if needed.
Initial DO cells will turn red if the value is 0.3 mg/L greater than the DO saturation point. If possible, reduce excess oxygen and prepare new sample.
Initial DO cells will turn blue if the value is 0.3 mg/L lower than the DO saturation point. If possible, increase oxygen and prepare new sample.
Code cells will turn green if a dilution of >200 mL is made; record extra nutrient addition.
Code cells will turn orange if a predilution must be made; record all dilution volumes.

## Pressure Verification
Monthly DO Meter Barometric Pressure Verification Log [version 6/30/23]
Facility Name
Meter or Barometer Model
Facility Elevation |  |  |  |  | feet
Month | Date | Initials | Official Source (that current local pressure was taken from) | Official Source BP (inches Hg) | Measured BP at Facility (mm Hg) | Official BP (mm Hg) | Local Elevation Correction Factor | Official Source Elevation Corrected BP (mm Hg) | Difference between facility and official BP (mm Hg) | Facility BP Adjustment Required? (adjust if >5 mm Hg difference)
JAN | [6 empty] | Enter elev'n above
FEB | [6 empty] | Enter elev'n above
MAR | [6 empty] | Enter elev'n above
APR | [6 empty] | Enter elev'n above
MAY | [6 empty] | Enter elev'n above
JUN | [6 empty] | Enter elev'n above
JUL | [6 empty] | Enter elev'n above
AUG | [6 empty] | Enter elev'n above
SEP | [6 empty] | Enter elev'n above
OCT | [6 empty] | Enter elev'n above
NOV | [6 empty] | Enter elev'n above
DEC | [6 empty] | Enter elev'n above
Comments:
PW=nr149
Local elevation correction = 760 mm Hg – (facility elevation in ft X 0.0254)
760 mm Hg
Example: Facility elevation is 850 feet above sea level
760 – (850 X 0.025)/760 = 0.972
Multiply official sea level corrected barometric pressure x 0.972 = elevation corrected barometric pressure
Note: To convert inches of Hg to mm of Hg multiply inches by 25.4
Example: 29.2 in Hg X 25.4 = 742 mm Hg
BP = Barometric Pressure
Barometric pressure must be verified at least annually, but it is recommended to verify monthly.

## Oxygen Sat. Table
|  | oC |  | Oxygen Saturation Chart | [14 empty] | y = mx +b |  | 759 | 18 |  | 9.4176
|  | in. Hg | 29.9213 | 29.7244 | 29.5276 | 29.3307 | 29.1339 | 28.937 | 28.7402 | 28.5433 | 28.3465 | 28.1496 | 27.9528 | 27.7559 | 27.5591 | 27.3622 | 27.1654 |  |  |  | 29.92 |  |  |  |  | 29.72 |  |  |  |  | 29.53 |  |  |  |  | 29.33 |  |  |  |  | 29.13 |  |  |  |  | 28.94 |  |  |  |  | 28.74 |  |  |  |  | 28.54 |  |  |  |  | 28.35 |  |  |  |  | 28.15 |  |  |  |  | 27.95 |  |  |  |  | 27.76 |  |  |  |  | 27.56 |  |  |  |  | 27.36 |  |  |  |  | 27.17
|  | mm Hg | 760 | 755 | 750 | 745 | 740 | 735 | 730 | 725 | 720 | 715 | 710 | 705 | 700 | 695 | 690 | m | b |  | 760 | 759 | 758 | 757 | 756 | 755 | 754 | 753 | 752 | 751 | 750 | 749 | 748 | 747 | 746 | 745 | 744 | 743 | 742 | 741 | 740 | 739 | 738 | 737 | 736 | 735 | 734 | 733 | 732 | 731 | 730 | 729 | 728 | 727 | 726 | 725 | 724 | 723 | 722 | 721 | 720 | 719 | 718 | 717 | 716 | 715 | 714 | 713 | 712 | 711 | 710 | 709 | 708 | 707 | 706 | 705 | 704 | 703 | 702 | 701 | 700 | 699 | 698 | 697 | 696 | 695 | 694 | 693 | 692 | 691 | 690
17 | 0 | 17 'C | 9.63 | 9.5666 | 9.5033 | 9.4399 | 9.3766 | 9.3132 | 9.2499 | 9.1865 | 9.1232 | 9.0598 | 8.9964 | 8.9331 | 8.8697 | 8.8064 | 8.743 | 0.012671 | 5.3291e-15 | 17 | 9.63 | 9.6173 | 9.6047 | 9.592 | 9.5793 | 9.5666 | 9.554 | 9.5413 | 9.5286 | 9.516 | 9.5033 | 9.4906 | 9.4779 | 9.4653 | 9.4526 | 9.4399 | 9.4273 | 9.4146 | 9.4019 | 9.3893 | 9.3766 | 9.3639 | 9.3512 | 9.3386 | 9.3259 | 9.3132 | 9.3006 | 9.2879 | 9.2752 | 9.2625 | 9.2499 | 9.2372 | 9.2245 | 9.2119 | 9.1992 | 9.1865 | 9.1738 | 9.1612 | 9.1485 | 9.1358 | 9.1232 | 9.1105 | 9.0978 | 9.0851 | 9.0725 | 9.0598 | 9.0471 | 9.0345 | 9.0218 | 9.0091 | 8.9964 | 8.9838 | 8.9711 | 8.9584 | 8.9458 | 8.9331 | 8.9204 | 8.9078 | 8.8951 | 8.8824 | 8.8697 | 8.8571 | 8.8444 | 8.8317 | 8.8191 | 8.8064 | 8.7937 | 8.781 | 8.7684 | 8.7557 | 8.743
17 | 0.1 | 17.1 'C | 9.61 | 9.5468 | 9.4836 | 9.4203 | 9.3571 | 9.2939 | 9.2307 | 9.1674 | 9.1042 | 9.041 | 8.9778 | 8.9145 | 8.8513 | 8.7881 | 8.7249 | 0.012645 | 1.7764e-15 | 17.1 | 9.61 | 9.5974 | 9.5847 | 9.5721 | 9.5594 | 9.5468 | 9.5341 | 9.5215 | 9.5088 | 9.4962 | 9.4836 | 9.4709 | 9.4583 | 9.4456 | 9.433 | 9.4203 | 9.4077 | 9.395 | 9.3824 | 9.3697 | 9.3571 | 9.3445 | 9.3318 | 9.3192 | 9.3065 | 9.2939 | 9.2812 | 9.2686 | 9.2559 | 9.2433 | 9.2307 | 9.218 | 9.2054 | 9.1927 | 9.1801 | 9.1674 | 9.1548 | 9.1421 | 9.1295 | 9.1169 | 9.1042 | 9.0916 | 9.0789 | 9.0663 | 9.0536 | 9.041 | 9.0283 | 9.0157 | 9.0031 | 8.9904 | 8.9778 | 8.9651 | 8.9525 | 8.9398 | 8.9272 | 8.9145 | 8.9019 | 8.8893 | 8.8766 | 8.864 | 8.8513 | 8.8387 | 8.826 | 8.8134 | 8.8007 | 8.7881 | 8.7754 | 8.7628 | 8.7502 | 8.7375 | 8.7249
17 | 0.2 | 17.2 'C | 9.59 | 9.5269 | 9.4638 | 9.4007 | 9.3376 | 9.2745 | 9.2114 | 9.1484 | 9.0853 | 9.0222 | 8.9591 | 8.896 | 8.8329 | 8.7698 | 8.7067 | 0.012618 | 1.2434e-14 | 17.2 | 9.59 | 9.5774 | 9.5648 | 9.5521 | 9.5395 | 9.5269 | 9.5143 | 9.5017 | 9.4891 | 9.4764 | 9.4638 | 9.4512 | 9.4386 | 9.426 | 9.4133 | 9.4007 | 9.3881 | 9.3755 | 9.3629 | 9.3503 | 9.3376 | 9.325 | 9.3124 | 9.2998 | 9.2872 | 9.2745 | 9.2619 | 9.2493 | 9.2367 | 9.2241 | 9.2114 | 9.1988 | 9.1862 | 9.1736 | 9.161 | 9.1484 | 9.1357 | 9.1231 | 9.1105 | 9.0979 | 9.0853 | 9.0726 | 9.06 | 9.0474 | 9.0348 | 9.0222 | 9.0096 | 8.9969 | 8.9843 | 8.9717 | 8.9591 | 8.9465 | 8.9338 | 8.9212 | 8.9086 | 8.896 | 8.8834 | 8.8707 | 8.8581 | 8.8455 | 8.8329 | 8.8203 | 8.8077 | 8.795 | 8.7824 | 8.7698 | 8.7572 | 8.7446 | 8.7319 | 8.7193 | 8.7067
17 | 0.3 | 17.3 'C | 9.57 | 9.507 | 9.4441 | 9.3811 | 9.3182 | 9.2552 | 9.1922 | 9.1293 | 9.0663 | 9.0034 | 8.9404 | 8.8774 | 8.8145 | 8.7515 | 8.6886 | 0.012592 | 3.5527e-15 | 17.3 | 9.57 | 9.5574 | 9.5448 | 9.5322 | 9.5196 | 9.507 | 9.4944 | 9.4819 | 9.4693 | 9.4567 | 9.4441 | 9.4315 | 9.4189 | 9.4063 | 9.3937 | 9.3811 | 9.3685 | 9.3559 | 9.3433 | 9.3308 | 9.3182 | 9.3056 | 9.293 | 9.2804 | 9.2678 | 9.2552 | 9.2426 | 9.23 | 9.2174 | 9.2048 | 9.1922 | 9.1796 | 9.1671 | 9.1545 | 9.1419 | 9.1293 | 9.1167 | 9.1041 | 9.0915 | 9.0789 | 9.0663 | 9.0537 | 9.0411 | 9.0285 | 9.0159 | 9.0034 | 8.9908 | 8.9782 | 8.9656 | 8.953 | 8.9404 | 8.9278 | 8.9152 | 8.9026 | 8.89 | 8.8774 | 8.8648 | 8.8522 | 8.8397 | 8.8271 | 8.8145 | 8.8019 | 8.7893 | 8.7767 | 8.7641 | 8.7515 | 8.7389 | 8.7263 | 8.7137 | 8.7011 | 8.6886
17 | 0.4 | 17.4 'C | 9.55 | 9.4872 | 9.4243 | 9.3615 | 9.2987 | 9.2359 | 9.173 | 9.1102 | 9.0474 | 8.9845 | 8.9217 | 8.8589 | 8.7961 | 8.7332 | 8.6704 | 0.012566 | 8.8818e-15 | 17.4 | 9.55 | 9.5374 | 9.5249 | 9.5123 | 9.4997 | 9.4872 | 9.4746 | 9.462 | 9.4495 | 9.4369 | 9.4243 | 9.4118 | 9.3992 | 9.3866 | 9.3741 | 9.3615 | 9.3489 | 9.3364 | 9.3238 | 9.3112 | 9.2987 | 9.2861 | 9.2736 | 9.261 | 9.2484 | 9.2359 | 9.2233 | 9.2107 | 9.1982 | 9.1856 | 9.173 | 9.1605 | 9.1479 | 9.1353 | 9.1228 | 9.1102 | 9.0976 | 9.0851 | 9.0725 | 9.0599 | 9.0474 | 9.0348 | 9.0222 | 9.0097 | 8.9971 | 8.9845 | 8.972 | 8.9594 | 8.9468 | 8.9343 | 8.9217 | 8.9091 | 8.8966 | 8.884 | 8.8714 | 8.8589 | 8.8463 | 8.8338 | 8.8212 | 8.8086 | 8.7961 | 8.7835 | 8.7709 | 8.7584 | 8.7458 | 8.7332 | 8.7207 | 8.7081 | 8.6955 | 8.683 | 8.6704
17 | 0.5 | 17.5 'C | 9.53 | 9.4673 | 9.4046 | 9.3419 | 9.2792 | 9.2165 | 9.1538 | 9.0911 | 9.0284 | 8.9657 | 8.903 | 8.8403 | 8.7776 | 8.7149 | 8.6522 | 0.012539 | 1.7764e-15 | 17.5 | 9.53 | 9.5175 | 9.5049 | 9.4924 | 9.4798 | 9.4673 | 9.4548 | 9.4422 | 9.4297 | 9.4171 | 9.4046 | 9.3921 | 9.3795 | 9.367 | 9.3544 | 9.3419 | 9.3294 | 9.3168 | 9.3043 | 9.2918 | 9.2792 | 9.2667 | 9.2541 | 9.2416 | 9.2291 | 9.2165 | 9.204 | 9.1914 | 9.1789 | 9.1664 | 9.1538 | 9.1413 | 9.1287 | 9.1162 | 9.1037 | 9.0911 | 9.0786 | 9.066 | 9.0535 | 9.041 | 9.0284 | 9.0159 | 9.0033 | 8.9908 | 8.9783 | 8.9657 | 8.9532 | 8.9406 | 8.9281 | 8.9156 | 8.903 | 8.8905 | 8.8779 | 8.8654 | 8.8529 | 8.8403 | 8.8278 | 8.8153 | 8.8027 | 8.7902 | 8.7776 | 8.7651 | 8.7526 | 8.74 | 8.7275 | 8.7149 | 8.7024 | 8.6899 | 8.6773 | 8.6648 | 8.6522
17 | 0.6 | 17.6 'C | 9.51 | 9.4474 | 9.3849 | 9.3223 | 9.2597 | 9.1972 | 9.1346 | 9.072 | 9.0095 | 8.9469 | 8.8843 | 8.8218 | 8.7592 | 8.6966 | 8.6341 | 0.012513 | -3.5527e-15 | 17.6 | 9.51 | 9.4975 | 9.485 | 9.4725 | 9.4599 | 9.4474 | 9.4349 | 9.4224 | 9.4099 | 9.3974 | 9.3849 | 9.3724 | 9.3598 | 9.3473 | 9.3348 | 9.3223 | 9.3098 | 9.2973 | 9.2848 | 9.2722 | 9.2597 | 9.2472 | 9.2347 | 9.2222 | 9.2097 | 9.1972 | 9.1847 | 9.1721 | 9.1596 | 9.1471 | 9.1346 | 9.1221 | 9.1096 | 9.0971 | 9.0846 | 9.072 | 9.0595 | 9.047 | 9.0345 | 9.022 | 9.0095 | 8.997 | 8.9844 | 8.9719 | 8.9594 | 8.9469 | 8.9344 | 8.9219 | 8.9094 | 8.8969 | 8.8843 | 8.8718 | 8.8593 | 8.8468 | 8.8343 | 8.8218 | 8.8093 | 8.7967 | 8.7842 | 8.7717 | 8.7592 | 8.7467 | 8.7342 | 8.7217 | 8.7092 | 8.6966 | 8.6841 | 8.6716 | 8.6591 | 8.6466 | 8.6341
17 | 0.7 | 17.7 'C | 9.49 | 9.4276 | 9.3651 | 9.3027 | 9.2403 | 9.1778 | 9.1154 | 9.053 | 8.9905 | 8.9281 | 8.8657 | 8.8032 | 8.7408 | 8.6784 | 8.6159 | 0.012487 | 7.1054e-15 | 17.7 | 9.49 | 9.4775 | 9.465 | 9.4525 | 9.4401 | 9.4276 | 9.4151 | 9.4026 | 9.3901 | 9.3776 | 9.3651 | 9.3526 | 9.3402 | 9.3277 | 9.3152 | 9.3027 | 9.2902 | 9.2777 | 9.2652 | 9.2528 | 9.2403 | 9.2278 | 9.2153 | 9.2028 | 9.1903 | 9.1778 | 9.1653 | 9.1529 | 9.1404 | 9.1279 | 9.1154 | 9.1029 | 9.0904 | 9.0779 | 9.0654 | 9.053 | 9.0405 | 9.028 | 9.0155 | 9.003 | 8.9905 | 8.978 | 8.9656 | 8.9531 | 8.9406 | 8.9281 | 8.9156 | 8.9031 | 8.8906 | 8.8781 | 8.8657 | 8.8532 | 8.8407 | 8.8282 | 8.8157 | 8.8032 | 8.7907 | 8.7782 | 8.7658 | 8.7533 | 8.7408 | 8.7283 | 8.7158 | 8.7033 | 8.6908 | 8.6784 | 8.6659 | 8.6534 | 8.6409 | 8.6284 | 8.6159
17 | 0.8 | 17.8 'C | 9.47 | 9.4077 | 9.3454 | 9.2831 | 9.2208 | 9.1585 | 9.0962 | 9.0339 | 8.9716 | 8.9093 | 8.847 | 8.7847 | 8.7224 | 8.6601 | 8.5978 | 0.012461 | -1.7764e-15 | 17.8 | 9.47 | 9.4575 | 9.4451 | 9.4326 | 9.4202 | 9.4077 | 9.3952 | 9.3828 | 9.3703 | 9.3579 | 9.3454 | 9.3329 | 9.3205 | 9.308 | 9.2956 | 9.2831 | 9.2706 | 9.2582 | 9.2457 | 9.2332 | 9.2208 | 9.2083 | 9.1959 | 9.1834 | 9.1709 | 9.1585 | 9.146 | 9.1336 | 9.1211 | 9.1086 | 9.0962 | 9.0837 | 9.0713 | 9.0588 | 9.0463 | 9.0339 | 9.0214 | 9.009 | 8.9965 | 8.984 | 8.9716 | 8.9591 | 8.9467 | 8.9342 | 8.9217 | 8.9093 | 8.8968 | 8.8844 | 8.8719 | 8.8594 | 8.847 | 8.8345 | 8.8221 | 8.8096 | 8.7971 | 8.7847 | 8.7722 | 8.7598 | 8.7473 | 8.7348 | 8.7224 | 8.7099 | 8.6974 | 8.685 | 8.6725 | 8.6601 | 8.6476 | 8.6351 | 8.6227 | 8.6102 | 8.5978
17 | 0.9 | 17.9 'C | 9.45 | 9.3878 | 9.3257 | 9.2635 | 9.2013 | 9.1391 | 9.077 | 9.0148 | 8.9526 | 8.8905 | 8.8283 | 8.7661 | 8.7039 | 8.6418 | 8.5796 | 0.012434 | 0 | 17.9 | 9.45 | 9.4376 | 9.4251 | 9.4127 | 9.4003 | 9.3878 | 9.3754 | 9.363 | 9.3505 | 9.3381 | 9.3257 | 9.3132 | 9.3008 | 9.2884 | 9.2759 | 9.2635 | 9.2511 | 9.2386 | 9.2262 | 9.2137 | 9.2013 | 9.1889 | 9.1764 | 9.164 | 9.1516 | 9.1391 | 9.1267 | 9.1143 | 9.1018 | 9.0894 | 9.077 | 9.0645 | 9.0521 | 9.0397 | 9.0272 | 9.0148 | 9.0024 | 8.9899 | 8.9775 | 8.9651 | 8.9526 | 8.9402 | 8.9278 | 8.9153 | 8.9029 | 8.8905 | 8.878 | 8.8656 | 8.8532 | 8.8407 | 8.8283 | 8.8159 | 8.8034 | 8.791 | 8.7786 | 8.7661 | 8.7537 | 8.7413 | 8.7288 | 8.7164 | 8.7039 | 8.6915 | 8.6791 | 8.6666 | 8.6542 | 8.6418 | 8.6293 | 8.6169 | 8.6045 | 8.592 | 8.5796
18 | 0 | 18 'C | 9.43 | 9.368 | 9.3059 | 9.2439 | 9.1818 | 9.1198 | 9.0578 | 8.9957 | 8.9337 | 8.8716 | 8.8096 | 8.7476 | 8.6855 | 8.6235 | 8.5614 | 0.012408 | 7.1054e-15 | 18 | 9.43 | 9.4176 | 9.4052 | 9.3928 | 9.3804 | 9.368 | 9.3556 | 9.3431 | 9.3307 | 9.3183 | 9.3059 | 9.2935 | 9.2811 | 9.2687 | 9.2563 | 9.2439 | 9.2315 | 9.2191 | 9.2067 | 9.1943 | 9.1818 | 9.1694 | 9.157 | 9.1446 | 9.1322 | 9.1198 | 9.1074 | 9.095 | 9.0826 | 9.0702 | 9.0578 | 9.0454 | 9.0329 | 9.0205 | 9.0081 | 8.9957 | 8.9833 | 8.9709 | 8.9585 | 8.9461 | 8.9337 | 8.9213 | 8.9089 | 8.8965 | 8.8841 | 8.8716 | 8.8592 | 8.8468 | 8.8344 | 8.822 | 8.8096 | 8.7972 | 8.7848 | 8.7724 | 8.76 | 8.7476 | 8.7352 | 8.7227 | 8.7103 | 8.6979 | 8.6855 | 8.6731 | 8.6607 | 8.6483 | 8.6359 | 8.6235 | 8.6111 | 8.5987 | 8.5863 | 8.5739 | 8.5614
18 | 0.1 | 18.1 'C | 9.41 | 9.3481 | 9.2862 | 9.2243 | 9.1624 | 9.1005 | 9.0386 | 8.9766 | 8.9147 | 8.8528 | 8.7909 | 8.729 | 8.6671 | 8.6052 | 8.5433 | 0.012382 | 8.8818e-15 | 18.1 | 9.41 | 9.3976 | 9.3852 | 9.3729 | 9.3605 | 9.3481 | 9.3357 | 9.3233 | 9.3109 | 9.2986 | 9.2862 | 9.2738 | 9.2614 | 9.249 | 9.2367 | 9.2243 | 9.2119 | 9.1995 | 9.1871 | 9.1747 | 9.1624 | 9.15 | 9.1376 | 9.1252 | 9.1128 | 9.1005 | 9.0881 | 9.0757 | 9.0633 | 9.0509 | 9.0386 | 9.0262 | 9.0138 | 9.0014 | 8.989 | 8.9766 | 8.9643 | 8.9519 | 8.9395 | 8.9271 | 8.9147 | 8.9024 | 8.89 | 8.8776 | 8.8652 | 8.8528 | 8.8404 | 8.8281 | 8.8157 | 8.8033 | 8.7909 | 8.7785 | 8.7662 | 8.7538 | 8.7414 | 8.729 | 8.7166 | 8.7043 | 8.6919 | 8.6795 | 8.6671 | 8.6547 | 8.6423 | 8.63 | 8.6176 | 8.6052 | 8.5928 | 8.5804 | 8.5681 | 8.5557 | 8.5433
18 | 0.2 | 18.2 'C | 9.39 | 9.3282 | 9.2664 | 9.2047 | 9.1429 | 9.0811 | 9.0193 | 8.9576 | 8.8958 | 8.834 | 8.7722 | 8.7105 | 8.6487 | 8.5869 | 8.5251 | 0.012355 | 0 | 18.2 | 9.39 | 9.3776 | 9.3653 | 9.3529 | 9.3406 | 9.3282 | 9.3159 | 9.3035 | 9.2912 | 9.2788 | 9.2664 | 9.2541 | 9.2417 | 9.2294 | 9.217 | 9.2047 | 9.1923 | 9.18 | 9.1676 | 9.1553 | 9.1429 | 9.1305 | 9.1182 | 9.1058 | 9.0935 | 9.0811 | 9.0688 | 9.0564 | 9.0441 | 9.0317 | 9.0193 | 9.007 | 8.9946 | 8.9823 | 8.9699 | 8.9576 | 8.9452 | 8.9329 | 8.9205 | 8.9081 | 8.8958 | 8.8834 | 8.8711 | 8.8587 | 8.8464 | 8.834 | 8.8217 | 8.8093 | 8.7969 | 8.7846 | 8.7722 | 8.7599 | 8.7475 | 8.7352 | 8.7228 | 8.7105 | 8.6981 | 8.6858 | 8.6734 | 8.661 | 8.6487 | 8.6363 | 8.624 | 8.6116 | 8.5993 | 8.5869 | 8.5746 | 8.5622 | 8.5498 | 8.5375 | 8.5251
18 | 0.3 | 18.3 'C | 9.37 | 9.3084 | 9.2467 | 9.1851 | 9.1234 | 9.0618 | 9.0001 | 8.9385 | 8.8768 | 8.8152 | 8.7536 | 8.6919 | 8.6303 | 8.5686 | 8.507 | 0.012329 | 8.8818e-15 | 18.3 | 9.37 | 9.3577 | 9.3453 | 9.333 | 9.3207 | 9.3084 | 9.296 | 9.2837 | 9.2714 | 9.259 | 9.2467 | 9.2344 | 9.2221 | 9.2097 | 9.1974 | 9.1851 | 9.1727 | 9.1604 | 9.1481 | 9.1357 | 9.1234 | 9.1111 | 9.0988 | 9.0864 | 9.0741 | 9.0618 | 9.0494 | 9.0371 | 9.0248 | 9.0125 | 9.0001 | 8.9878 | 8.9755 | 8.9631 | 8.9508 | 8.9385 | 8.9262 | 8.9138 | 8.9015 | 8.8892 | 8.8768 | 8.8645 | 8.8522 | 8.8399 | 8.8275 | 8.8152 | 8.8029 | 8.7905 | 8.7782 | 8.7659 | 8.7536 | 8.7412 | 8.7289 | 8.7166 | 8.7042 | 8.6919 | 8.6796 | 8.6672 | 8.6549 | 8.6426 | 8.6303 | 8.6179 | 8.6056 | 8.5933 | 8.5809 | 8.5686 | 8.5563 | 8.544 | 8.5316 | 8.5193 | 8.507
18 | 0.4 | 18.4 'C | 9.35 | 9.2885 | 9.227 | 9.1655 | 9.1039 | 9.0424 | 8.9809 | 8.9194 | 8.8579 | 8.7964 | 8.7349 | 8.6734 | 8.6118 | 8.5503 | 8.4888 | 0.012303 | 3.5527e-15 | 18.4 | 9.35 | 9.3377 | 9.3254 | 9.3131 | 9.3008 | 9.2885 | 9.2762 | 9.2639 | 9.2516 | 9.2393 | 9.227 | 9.2147 | 9.2024 | 9.1901 | 9.1778 | 9.1655 | 9.1532 | 9.1409 | 9.1286 | 9.1163 | 9.1039 | 9.0916 | 9.0793 | 9.067 | 9.0547 | 9.0424 | 9.0301 | 9.0178 | 9.0055 | 8.9932 | 8.9809 | 8.9686 | 8.9563 | 8.944 | 8.9317 | 8.9194 | 8.9071 | 8.8948 | 8.8825 | 8.8702 | 8.8579 | 8.8456 | 8.8333 | 8.821 | 8.8087 | 8.7964 | 8.7841 | 8.7718 | 8.7595 | 8.7472 | 8.7349 | 8.7226 | 8.7103 | 8.698 | 8.6857 | 8.6734 | 8.6611 | 8.6487 | 8.6364 | 8.6241 | 8.6118 | 8.5995 | 8.5872 | 8.5749 | 8.5626 | 8.5503 | 8.538 | 8.5257 | 8.5134 | 8.5011 | 8.4888
18 | 0.5 | 18.5 'C | 9.34 | 9.2786 | 9.2171 | 9.1557 | 9.0942 | 9.0328 | 8.9713 | 8.9099 | 8.8484 | 8.787 | 8.7255 | 8.6641 | 8.6026 | 8.5412 | 8.4797 | 0.012289 | 5.3291e-15 | 18.5 | 9.34 | 9.3277 | 9.3154 | 9.3031 | 9.2908 | 9.2786 | 9.2663 | 9.254 | 9.2417 | 9.2294 | 9.2171 | 9.2048 | 9.1925 | 9.1802 | 9.1679 | 9.1557 | 9.1434 | 9.1311 | 9.1188 | 9.1065 | 9.0942 | 9.0819 | 9.0696 | 9.0573 | 9.0451 | 9.0328 | 9.0205 | 9.0082 | 8.9959 | 8.9836 | 8.9713 | 8.959 | 8.9467 | 8.9344 | 8.9222 | 8.9099 | 8.8976 | 8.8853 | 8.873 | 8.8607 | 8.8484 | 8.8361 | 8.8238 | 8.8116 | 8.7993 | 8.787 | 8.7747 | 8.7624 | 8.7501 | 8.7378 | 8.7255 | 8.7132 | 8.7009 | 8.6887 | 8.6764 | 8.6641 | 8.6518 | 8.6395 | 8.6272 | 8.6149 | 8.6026 | 8.5903 | 8.5781 | 8.5658 | 8.5535 | 8.5412 | 8.5289 | 8.5166 | 8.5043 | 8.492 | 8.4797
18 | 0.6 | 18.6 'C | 9.32 | 9.2587 | 9.1974 | 9.1361 | 9.0747 | 9.0134 | 8.9521 | 8.8908 | 8.8295 | 8.7682 | 8.7068 | 8.6455 | 8.5842 | 8.5229 | 8.4616 | 0.012263 | 3.5527e-15 | 18.6 | 9.32 | 9.3077 | 9.2955 | 9.2832 | 9.2709 | 9.2587 | 9.2464 | 9.2342 | 9.2219 | 9.2096 | 9.1974 | 9.1851 | 9.1728 | 9.1606 | 9.1483 | 9.1361 | 9.1238 | 9.1115 | 9.0993 | 9.087 | 9.0747 | 9.0625 | 9.0502 | 9.0379 | 9.0257 | 9.0134 | 9.0012 | 8.9889 | 8.9766 | 8.9644 | 8.9521 | 8.9398 | 8.9276 | 8.9153 | 8.9031 | 8.8908 | 8.8785 | 8.8663 | 8.854 | 8.8417 | 8.8295 | 8.8172 | 8.8049 | 8.7927 | 8.7804 | 8.7682 | 8.7559 | 8.7436 | 8.7314 | 8.7191 | 8.7068 | 8.6946 | 8.6823 | 8.6701 | 8.6578 | 8.6455 | 8.6333 | 8.621 | 8.6087 | 8.5965 | 8.5842 | 8.5719 | 8.5597 | 8.5474 | 8.5352 | 8.5229 | 8.5106 | 8.4984 | 8.4861 | 8.4738 | 8.4616
18 | 0.7 | 18.7 'C | 9.3 | 9.2388 | 9.1776 | 9.1164 | 9.0553 | 8.9941 | 8.9329 | 8.8717 | 8.8105 | 8.7493 | 8.6882 | 8.627 | 8.5658 | 8.5046 | 8.4434 | 0.012237 | 1.7764e-15 | 18.7 | 9.3 | 9.2878 | 9.2755 | 9.2633 | 9.2511 | 9.2388 | 9.2266 | 9.2143 | 9.2021 | 9.1899 | 9.1776 | 9.1654 | 9.1532 | 9.1409 | 9.1287 | 9.1164 | 9.1042 | 9.092 | 9.0797 | 9.0675 | 9.0553 | 9.043 | 9.0308 | 9.0186 | 9.0063 | 8.9941 | 8.9818 | 8.9696 | 8.9574 | 8.9451 | 8.9329 | 8.9207 | 8.9084 | 8.8962 | 8.8839 | 8.8717 | 8.8595 | 8.8472 | 8.835 | 8.8228 | 8.8105 | 8.7983 | 8.7861 | 8.7738 | 8.7616 | 8.7493 | 8.7371 | 8.7249 | 8.7126 | 8.7004 | 8.6882 | 8.6759 | 8.6637 | 8.6514 | 8.6392 | 8.627 | 8.6147 | 8.6025 | 8.5903 | 8.578 | 8.5658 | 8.5536 | 8.5413 | 8.5291 | 8.5168 | 8.5046 | 8.4924 | 8.4801 | 8.4679 | 8.4557 | 8.4434
18 | 0.8 | 18.8 'C | 9.28 | 9.2189 | 9.1579 | 9.0968 | 9.0358 | 8.9747 | 8.9137 | 8.8526 | 8.7916 | 8.7305 | 8.6695 | 8.6084 | 8.5474 | 8.4863 | 8.4253 | 0.012211 | 8.8818e-15 | 18.8 | 9.28 | 9.2678 | 9.2556 | 9.2434 | 9.2312 | 9.2189 | 9.2067 | 9.1945 | 9.1823 | 9.1701 | 9.1579 | 9.1457 | 9.1335 | 9.1213 | 9.1091 | 9.0968 | 9.0846 | 9.0724 | 9.0602 | 9.048 | 9.0358 | 9.0236 | 9.0114 | 8.9992 | 8.9869 | 8.9747 | 8.9625 | 8.9503 | 8.9381 | 8.9259 | 8.9137 | 8.9015 | 8.8893 | 8.8771 | 8.8648 | 8.8526 | 8.8404 | 8.8282 | 8.816 | 8.8038 | 8.7916 | 8.7794 | 8.7672 | 8.7549 | 8.7427 | 8.7305 | 8.7183 | 8.7061 | 8.6939 | 8.6817 | 8.6695 | 8.6573 | 8.6451 | 8.6328 | 8.6206 | 8.6084 | 8.5962 | 8.584 | 8.5718 | 8.5596 | 8.5474 | 8.5352 | 8.5229 | 8.5107 | 8.4985 | 8.4863 | 8.4741 | 8.4619 | 8.4497 | 8.4375 | 8.4253
18 | 0.9 | 18.9 'C | 9.26 | 9.1991 | 9.1382 | 9.0772 | 9.0163 | 8.9554 | 8.8945 | 8.8336 | 8.7726 | 8.7117 | 8.6508 | 8.5899 | 8.5289 | 8.468 | 8.4071 | 0.012184 | 0 | 18.9 | 9.26 | 9.2478 | 9.2356 | 9.2234 | 9.2113 | 9.1991 | 9.1869 | 9.1747 | 9.1625 | 9.1503 | 9.1382 | 9.126 | 9.1138 | 9.1016 | 9.0894 | 9.0772 | 9.0651 | 9.0529 | 9.0407 | 9.0285 | 9.0163 | 9.0041 | 8.9919 | 8.9798 | 8.9676 | 8.9554 | 8.9432 | 8.931 | 8.9188 | 8.9067 | 8.8945 | 8.8823 | 8.8701 | 8.8579 | 8.8457 | 8.8336 | 8.8214 | 8.8092 | 8.797 | 8.7848 | 8.7726 | 8.7604 | 8.7483 | 8.7361 | 8.7239 | 8.7117 | 8.6995 | 8.6873 | 8.6752 | 8.663 | 8.6508 | 8.6386 | 8.6264 | 8.6142 | 8.6021 | 8.5899 | 8.5777 | 8.5655 | 8.5533 | 8.5411 | 8.5289 | 8.5168 | 8.5046 | 8.4924 | 8.4802 | 8.468 | 8.4558 | 8.4437 | 8.4315 | 8.4193 | 8.4071
19 | 0 | 19 'C | 9.24 | 9.1792 | 9.1184 | 9.0576 | 8.9968 | 8.9361 | 8.8753 | 8.8145 | 8.7537 | 8.6929 | 8.6321 | 8.5713 | 8.5105 | 8.4497 | 8.3889 | 0.012158 | -1.7764e-15 | 19 | 9.24 | 9.2278 | 9.2157 | 9.2035 | 9.1914 | 9.1792 | 9.1671 | 9.1549 | 9.1427 | 9.1306 | 9.1184 | 9.1063 | 9.0941 | 9.0819 | 9.0698 | 9.0576 | 9.0455 | 9.0333 | 9.0212 | 9.009 | 8.9968 | 8.9847 | 8.9725 | 8.9604 | 8.9482 | 8.9361 | 8.9239 | 8.9117 | 8.8996 | 8.8874 | 8.8753 | 8.8631 | 8.8509 | 8.8388 | 8.8266 | 8.8145 | 8.8023 | 8.7902 | 8.778 | 8.7658 | 8.7537 | 8.7415 | 8.7294 | 8.7172 | 8.7051 | 8.6929 | 8.6807 | 8.6686 | 8.6564 | 8.6443 | 8.6321 | 8.6199 | 8.6078 | 8.5956 | 8.5835 | 8.5713 | 8.5592 | 8.547 | 8.5348 | 8.5227 | 8.5105 | 8.4984 | 8.4862 | 8.4741 | 8.4619 | 8.4497 | 8.4376 | 8.4254 | 8.4133 | 8.4011 | 8.3889
19 | 0.1 | 19.1 'C | 9.22 | 9.1593 | 9.0987 | 9.038 | 8.9774 | 8.9167 | 8.8561 | 8.7954 | 8.7347 | 8.6741 | 8.6134 | 8.5528 | 8.4921 | 8.4314 | 8.3708 | 0.012132 | 5.3291e-15 | 19.1 | 9.22 | 9.2079 | 9.1957 | 9.1836 | 9.1715 | 9.1593 | 9.1472 | 9.1351 | 9.1229 | 9.1108 | 9.0987 | 9.0866 | 9.0744 | 9.0623 | 9.0502 | 9.038 | 9.0259 | 9.0138 | 9.0016 | 8.9895 | 8.9774 | 8.9652 | 8.9531 | 8.941 | 8.9288 | 8.9167 | 8.9046 | 8.8924 | 8.8803 | 8.8682 | 8.8561 | 8.8439 | 8.8318 | 8.8197 | 8.8075 | 8.7954 | 8.7833 | 8.7711 | 8.759 | 8.7469 | 8.7347 | 8.7226 | 8.7105 | 8.6983 | 8.6862 | 8.6741 | 8.6619 | 8.6498 | 8.6377 | 8.6256 | 8.6134 | 8.6013 | 8.5892 | 8.577 | 8.5649 | 8.5528 | 8.5406 | 8.5285 | 8.5164 | 8.5042 | 8.4921 | 8.48 | 8.4678 | 8.4557 | 8.4436 | 8.4314 | 8.4193 | 8.4072 | 8.3951 | 8.3829 | 8.3708
19 | 0.2 | 19.2 'C | 9.2 | 9.1395 | 9.0789 | 9.0184 | 8.9579 | 8.8974 | 8.8368 | 8.7763 | 8.7158 | 8.6553 | 8.5947 | 8.5342 | 8.4737 | 8.4132 | 8.3526 | 0.012105 | 5.3291e-15 | 19.2 | 9.2 | 9.1879 | 9.1758 | 9.1637 | 9.1516 | 9.1395 | 9.1274 | 9.1153 | 9.1032 | 9.0911 | 9.0789 | 9.0668 | 9.0547 | 9.0426 | 9.0305 | 9.0184 | 9.0063 | 8.9942 | 8.9821 | 8.97 | 8.9579 | 8.9458 | 8.9337 | 8.9216 | 8.9095 | 8.8974 | 8.8853 | 8.8732 | 8.8611 | 8.8489 | 8.8368 | 8.8247 | 8.8126 | 8.8005 | 8.7884 | 8.7763 | 8.7642 | 8.7521 | 8.74 | 8.7279 | 8.7158 | 8.7037 | 8.6916 | 8.6795 | 8.6674 | 8.6553 | 8.6432 | 8.6311 | 8.6189 | 8.6068 | 8.5947 | 8.5826 | 8.5705 | 8.5584 | 8.5463 | 8.5342 | 8.5221 | 8.51 | 8.4979 | 8.4858 | 8.4737 | 8.4616 | 8.4495 | 8.4374 | 8.4253 | 8.4132 | 8.4011 | 8.3889 | 8.3768 | 8.3647 | 8.3526
19 | 0.3 | 19.3 'C | 9.19 | 9.1295 | 9.0691 | 9.0086 | 8.9482 | 8.8877 | 8.8272 | 8.7668 | 8.7063 | 8.6459 | 8.5854 | 8.5249 | 8.4645 | 8.404 | 8.3436 | 0.012092 | 1.4211e-14 | 19.3 | 9.19 | 9.1779 | 9.1658 | 9.1537 | 9.1416 | 9.1295 | 9.1174 | 9.1054 | 9.0933 | 9.0812 | 9.0691 | 9.057 | 9.0449 | 9.0328 | 9.0207 | 9.0086 | 8.9965 | 8.9844 | 8.9723 | 8.9603 | 8.9482 | 8.9361 | 8.924 | 8.9119 | 8.8998 | 8.8877 | 8.8756 | 8.8635 | 8.8514 | 8.8393 | 8.8272 | 8.8151 | 8.8031 | 8.791 | 8.7789 | 8.7668 | 8.7547 | 8.7426 | 8.7305 | 8.7184 | 8.7063 | 8.6942 | 8.6821 | 8.67 | 8.6579 | 8.6459 | 8.6338 | 8.6217 | 8.6096 | 8.5975 | 8.5854 | 8.5733 | 8.5612 | 8.5491 | 8.537 | 8.5249 | 8.5128 | 8.5008 | 8.4887 | 8.4766 | 8.4645 | 8.4524 | 8.4403 | 8.4282 | 8.4161 | 8.404 | 8.3919 | 8.3798 | 8.3677 | 8.3556 | 8.3436
19 | 0.4 | 19.4 'C | 9.17 | 9.1097 | 9.0493 | 8.989 | 8.9287 | 8.8684 | 8.808 | 8.7477 | 8.6874 | 8.627 | 8.5667 | 8.5064 | 8.4461 | 8.3857 | 8.3254 | 0.012066 | 1.7764e-15 | 19.4 | 9.17 | 9.1579 | 9.1459 | 9.1338 | 9.1217 | 9.1097 | 9.0976 | 9.0855 | 9.0735 | 9.0614 | 9.0493 | 9.0373 | 9.0252 | 9.0131 | 9.0011 | 8.989 | 8.9769 | 8.9649 | 8.9528 | 8.9407 | 8.9287 | 8.9166 | 8.9046 | 8.8925 | 8.8804 | 8.8684 | 8.8563 | 8.8442 | 8.8322 | 8.8201 | 8.808 | 8.796 | 8.7839 | 8.7718 | 8.7598 | 8.7477 | 8.7356 | 8.7236 | 8.7115 | 8.6994 | 8.6874 | 8.6753 | 8.6632 | 8.6512 | 8.6391 | 8.627 | 8.615 | 8.6029 | 8.5908 | 8.5788 | 8.5667 | 8.5546 | 8.5426 | 8.5305 | 8.5184 | 8.5064 | 8.4943 | 8.4823 | 8.4702 | 8.4581 | 8.4461 | 8.434 | 8.4219 | 8.4099 | 8.3978 | 8.3857 | 8.3737 | 8.3616 | 8.3495 | 8.3375 | 8.3254
19 | 0.5 | 19.5 'C | 9.15 | 9.0898 | 9.0296 | 8.9694 | 8.9092 | 8.849 | 8.7888 | 8.7286 | 8.6684 | 8.6082 | 8.548 | 8.4878 | 8.4276 | 8.3674 | 8.3072 | 0.012039 | -3.5527e-15 | 19.5 | 9.15 | 9.138 | 9.1259 | 9.1139 | 9.1018 | 9.0898 | 9.0778 | 9.0657 | 9.0537 | 9.0416 | 9.0296 | 9.0176 | 9.0055 | 8.9935 | 8.9814 | 8.9694 | 8.9574 | 8.9453 | 8.9333 | 8.9213 | 8.9092 | 8.8972 | 8.8851 | 8.8731 | 8.8611 | 8.849 | 8.837 | 8.8249 | 8.8129 | 8.8009 | 8.7888 | 8.7768 | 8.7647 | 8.7527 | 8.7407 | 8.7286 | 8.7166 | 8.7045 | 8.6925 | 8.6805 | 8.6684 | 8.6564 | 8.6443 | 8.6323 | 8.6203 | 8.6082 | 8.5962 | 8.5841 | 8.5721 | 8.5601 | 8.548 | 8.536 | 8.5239 | 8.5119 | 8.4999 | 8.4878 | 8.4758 | 8.4637 | 8.4517 | 8.4397 | 8.4276 | 8.4156 | 8.4036 | 8.3915 | 8.3795 | 8.3674 | 8.3554 | 8.3434 | 8.3313 | 8.3193 | 8.3072
19 | 0.6 | 19.6 'C | 9.13 | 9.0699 | 9.0099 | 8.9498 | 8.8897 | 8.8297 | 8.7696 | 8.7095 | 8.6495 | 8.5894 | 8.5293 | 8.4693 | 8.4092 | 8.3491 | 8.2891 | 0.012013 | 7.1054e-15 | 19.6 | 9.13 | 9.118 | 9.106 | 9.094 | 9.0819 | 9.0699 | 9.0579 | 9.0459 | 9.0339 | 9.0219 | 9.0099 | 8.9979 | 8.9858 | 8.9738 | 8.9618 | 8.9498 | 8.9378 | 8.9258 | 8.9138 | 8.9017 | 8.8897 | 8.8777 | 8.8657 | 8.8537 | 8.8417 | 8.8297 | 8.8177 | 8.8056 | 8.7936 | 8.7816 | 8.7696 | 8.7576 | 8.7456 | 8.7336 | 8.7216 | 8.7095 | 8.6975 | 8.6855 | 8.6735 | 8.6615 | 8.6495 | 8.6375 | 8.6254 | 8.6134 | 8.6014 | 8.5894 | 8.5774 | 8.5654 | 8.5534 | 8.5414 | 8.5293 | 8.5173 | 8.5053 | 8.4933 | 8.4813 | 8.4693 | 8.4573 | 8.4452 | 8.4332 | 8.4212 | 8.4092 | 8.3972 | 8.3852 | 8.3732 | 8.3612 | 8.3491 | 8.3371 | 8.3251 | 8.3131 | 8.3011 | 8.2891
19 | 0.7 | 19.7 'C | 9.11 | 9.0501 | 8.9901 | 8.9302 | 8.8703 | 8.8103 | 8.7504 | 8.6905 | 8.6305 | 8.5706 | 8.5107 | 8.4507 | 8.3908 | 8.3309 | 8.2709 | 0.011987 | -1.7764e-15 | 19.7 | 9.11 | 9.098 | 9.086 | 9.074 | 9.0621 | 9.0501 | 9.0381 | 9.0261 | 9.0141 | 9.0021 | 8.9901 | 8.9781 | 8.9662 | 8.9542 | 8.9422 | 8.9302 | 8.9182 | 8.9062 | 8.8942 | 8.8823 | 8.8703 | 8.8583 | 8.8463 | 8.8343 | 8.8223 | 8.8103 | 8.7983 | 8.7864 | 8.7744 | 8.7624 | 8.7504 | 8.7384 | 8.7264 | 8.7144 | 8.7024 | 8.6905 | 8.6785 | 8.6665 | 8.6545 | 8.6425 | 8.6305 | 8.6185 | 8.6066 | 8.5946 | 8.5826 | 8.5706 | 8.5586 | 8.5466 | 8.5346 | 8.5226 | 8.5107 | 8.4987 | 8.4867 | 8.4747 | 8.4627 | 8.4507 | 8.4387 | 8.4268 | 8.4148 | 8.4028 | 8.3908 | 8.3788 | 8.3668 | 8.3548 | 8.3428 | 8.3309 | 8.3189 | 8.3069 | 8.2949 | 8.2829 | 8.2709
19 | 0.8 | 19.8 'C | 9.1 | 9.0401 | 8.9803 | 8.9204 | 8.8605 | 8.8007 | 8.7408 | 8.6809 | 8.6211 | 8.5612 | 8.5013 | 8.4414 | 8.3816 | 8.3217 | 8.2618 | 0.011974 | -1.7764e-15 | 19.8 | 9.1 | 9.088 | 9.0761 | 9.0641 | 9.0521 | 9.0401 | 9.0282 | 9.0162 | 9.0042 | 8.9922 | 8.9803 | 8.9683 | 8.9563 | 8.9443 | 8.9324 | 8.9204 | 8.9084 | 8.8964 | 8.8845 | 8.8725 | 8.8605 | 8.8486 | 8.8366 | 8.8246 | 8.8126 | 8.8007 | 8.7887 | 8.7767 | 8.7647 | 8.7528 | 8.7408 | 8.7288 | 8.7168 | 8.7049 | 8.6929 | 8.6809 | 8.6689 | 8.657 | 8.645 | 8.633 | 8.6211 | 8.6091 | 8.5971 | 8.5851 | 8.5732 | 8.5612 | 8.5492 | 8.5372 | 8.5253 | 8.5133 | 8.5013 | 8.4893 | 8.4774 | 8.4654 | 8.4534 | 8.4414 | 8.4295 | 8.4175 | 8.4055 | 8.3936 | 8.3816 | 8.3696 | 8.3576 | 8.3457 | 8.3337 | 8.3217 | 8.3097 | 8.2978 | 8.2858 | 8.2738 | 8.2618
19 | 0.9 | 19.9 'C | 9.08 | 9.0203 | 8.9605 | 8.9008 | 8.8411 | 8.7813 | 8.7216 | 8.6618 | 8.6021 | 8.5424 | 8.4826 | 8.4229 | 8.3632 | 8.3034 | 8.2437 | 0.011947 | 3.5527e-15 | 19.9 | 9.08 | 9.0681 | 9.0561 | 9.0442 | 9.0322 | 9.0203 | 9.0083 | 8.9964 | 8.9844 | 8.9725 | 8.9605 | 8.9486 | 8.9366 | 8.9247 | 8.9127 | 8.9008 | 8.8888 | 8.8769 | 8.8649 | 8.853 | 8.8411 | 8.8291 | 8.8172 | 8.8052 | 8.7933 | 8.7813 | 8.7694 | 8.7574 | 8.7455 | 8.7335 | 8.7216 | 8.7096 | 8.6977 | 8.6857 | 8.6738 | 8.6618 | 8.6499 | 8.6379 | 8.626 | 8.6141 | 8.6021 | 8.5902 | 8.5782 | 8.5663 | 8.5543 | 8.5424 | 8.5304 | 8.5185 | 8.5065 | 8.4946 | 8.4826 | 8.4707 | 8.4587 | 8.4468 | 8.4348 | 8.4229 | 8.4109 | 8.399 | 8.3871 | 8.3751 | 8.3632 | 8.3512 | 8.3393 | 8.3273 | 8.3154 | 8.3034 | 8.2915 | 8.2795 | 8.2676 | 8.2556 | 8.2437
20 | 0 | 20 'C | 9.06 | 9.0004 | 8.9408 | 8.8812 | 8.8216 | 8.762 | 8.7024 | 8.6428 | 8.5832 | 8.5236 | 8.4639 | 8.4043 | 8.3447 | 8.2851 | 8.2255 | 0.011921 | 5.3291e-15 | 20 | 9.06 | 9.0481 | 9.0362 | 9.0242 | 9.0123 | 9.0004 | 8.9885 | 8.9766 | 8.9646 | 8.9527 | 8.9408 | 8.9289 | 8.9169 | 8.905 | 8.8931 | 8.8812 | 8.8693 | 8.8573 | 8.8454 | 8.8335 | 8.8216 | 8.8097 | 8.7977 | 8.7858 | 8.7739 | 8.762 | 8.7501 | 8.7381 | 8.7262 | 8.7143 | 8.7024 | 8.6904 | 8.6785 | 8.6666 | 8.6547 | 8.6428 | 8.6308 | 8.6189 | 8.607 | 8.5951 | 8.5832 | 8.5712 | 8.5593 | 8.5474 | 8.5355 | 8.5236 | 8.5116 | 8.4997 | 8.4878 | 8.4759 | 8.4639 | 8.452 | 8.4401 | 8.4282 | 8.4163 | 8.4043 | 8.3924 | 8.3805 | 8.3686 | 8.3567 | 8.3447 | 8.3328 | 8.3209 | 8.309 | 8.2971 | 8.2851 | 8.2732 | 8.2613 | 8.2494 | 8.2374 | 8.2255
20 | 0.1 | 20.1 'C | 9.04 | 8.9805 | 8.9211 | 8.8616 | 8.8021 | 8.7426 | 8.6832 | 8.6237 | 8.5642 | 8.5047 | 8.4453 | 8.3858 | 8.3263 | 8.2668 | 8.2074 | 0.011895 | 0 | 20.1 | 9.04 | 9.0281 | 9.0162 | 9.0043 | 8.9924 | 8.9805 | 8.9686 | 8.9567 | 8.9448 | 8.9329 | 8.9211 | 8.9092 | 8.8973 | 8.8854 | 8.8735 | 8.8616 | 8.8497 | 8.8378 | 8.8259 | 8.814 | 8.8021 | 8.7902 | 8.7783 | 8.7664 | 8.7545 | 8.7426 | 8.7307 | 8.7188 | 8.7069 | 8.6951 | 8.6832 | 8.6713 | 8.6594 | 8.6475 | 8.6356 | 8.6237 | 8.6118 | 8.5999 | 8.588 | 8.5761 | 8.5642 | 8.5523 | 8.5404 | 8.5285 | 8.5166 | 8.5047 | 8.4928 | 8.4809 | 8.4691 | 8.4572 | 8.4453 | 8.4334 | 8.4215 | 8.4096 | 8.3977 | 8.3858 | 8.3739 | 8.362 | 8.3501 | 8.3382 | 8.3263 | 8.3144 | 8.3025 | 8.2906 | 8.2787 | 8.2668 | 8.2549 | 8.2431 | 8.2312 | 8.2193 | 8.2074
20 | 0.2 | 20.2 'C | 9.02 | 8.9607 | 8.9013 | 8.842 | 8.7826 | 8.7233 | 8.6639 | 8.6046 | 8.5453 | 8.4859 | 8.4266 | 8.3672 | 8.3079 | 8.2486 | 8.1892 | 0.011868 | 7.1054e-15 | 20.2 | 9.02 | 9.0081 | 8.9963 | 8.9844 | 8.9725 | 8.9607 | 8.9488 | 8.9369 | 8.9251 | 8.9132 | 8.9013 | 8.8894 | 8.8776 | 8.8657 | 8.8538 | 8.842 | 8.8301 | 8.8182 | 8.8064 | 8.7945 | 8.7826 | 8.7708 | 8.7589 | 8.747 | 8.7352 | 8.7233 | 8.7114 | 8.6996 | 8.6877 | 8.6758 | 8.6639 | 8.6521 | 8.6402 | 8.6283 | 8.6165 | 8.6046 | 8.5927 | 8.5809 | 8.569 | 8.5571 | 8.5453 | 8.5334 | 8.5215 | 8.5097 | 8.4978 | 8.4859 | 8.4741 | 8.4622 | 8.4503 | 8.4384 | 8.4266 | 8.4147 | 8.4028 | 8.391 | 8.3791 | 8.3672 | 8.3554 | 8.3435 | 8.3316 | 8.3198 | 8.3079 | 8.296 | 8.2842 | 8.2723 | 8.2604 | 8.2486 | 8.2367 | 8.2248 | 8.2129 | 8.2011 | 8.1892
20 | 0.3 | 20.3 'C | 9.01 | 8.9507 | 8.8914 | 8.8322 | 8.7729 | 8.7136 | 8.6543 | 8.5951 | 8.5358 | 8.4765 | 8.4172 | 8.358 | 8.2987 | 8.2394 | 8.1801 | 0.011855 | 1.2434e-14 | 20.3 | 9.01 | 8.9981 | 8.9863 | 8.9744 | 8.9626 | 8.9507 | 8.9389 | 8.927 | 8.9152 | 8.9033 | 8.8914 | 8.8796 | 8.8677 | 8.8559 | 8.844 | 8.8322 | 8.8203 | 8.8085 | 8.7966 | 8.7848 | 8.7729 | 8.761 | 8.7492 | 8.7373 | 8.7255 | 8.7136 | 8.7018 | 8.6899 | 8.6781 | 8.6662 | 8.6543 | 8.6425 | 8.6306 | 8.6188 | 8.6069 | 8.5951 | 8.5832 | 8.5714 | 8.5595 | 8.5476 | 8.5358 | 8.5239 | 8.5121 | 8.5002 | 8.4884 | 8.4765 | 8.4647 | 8.4528 | 8.4409 | 8.4291 | 8.4172 | 8.4054 | 8.3935 | 8.3817 | 8.3698 | 8.358 | 8.3461 | 8.3343 | 8.3224 | 8.3105 | 8.2987 | 8.2868 | 8.275 | 8.2631 | 8.2513 | 8.2394 | 8.2276 | 8.2157 | 8.2038 | 8.192 | 8.1801
20 | 0.4 | 20.4 'C | 8.99 | 8.9309 | 8.8717 | 8.8126 | 8.7534 | 8.6943 | 8.6351 | 8.576 | 8.5168 | 8.4577 | 8.3986 | 8.3394 | 8.2803 | 8.2211 | 8.162 | 0.011829 | 3.5527e-15 | 20.4 | 8.99 | 8.9782 | 8.9663 | 8.9545 | 8.9427 | 8.9309 | 8.919 | 8.9072 | 8.8954 | 8.8835 | 8.8717 | 8.8599 | 8.8481 | 8.8362 | 8.8244 | 8.8126 | 8.8007 | 8.7889 | 8.7771 | 8.7652 | 8.7534 | 8.7416 | 8.7298 | 8.7179 | 8.7061 | 8.6943 | 8.6824 | 8.6706 | 8.6588 | 8.647 | 8.6351 | 8.6233 | 8.6115 | 8.5996 | 8.5878 | 8.576 | 8.5642 | 8.5523 | 8.5405 | 8.5287 | 8.5168 | 8.505 | 8.4932 | 8.4814 | 8.4695 | 8.4577 | 8.4459 | 8.434 | 8.4222 | 8.4104 | 8.3986 | 8.3867 | 8.3749 | 8.3631 | 8.3512 | 8.3394 | 8.3276 | 8.3157 | 8.3039 | 8.2921 | 8.2803 | 8.2684 | 8.2566 | 8.2448 | 8.2329 | 8.2211 | 8.2093 | 8.1975 | 8.1856 | 8.1738 | 8.162
20 | 0.5 | 20.5 'C | 8.97 | 8.911 | 8.852 | 8.793 | 8.7339 | 8.6749 | 8.6159 | 8.5569 | 8.4979 | 8.4389 | 8.3799 | 8.3209 | 8.2618 | 8.2028 | 8.1438 | 0.011803 | -5.3291e-15 | 20.5 | 8.97 | 8.9582 | 8.9464 | 8.9346 | 8.9228 | 8.911 | 8.8992 | 8.8874 | 8.8756 | 8.8638 | 8.852 | 8.8402 | 8.8284 | 8.8166 | 8.8048 | 8.793 | 8.7812 | 8.7694 | 8.7576 | 8.7457 | 8.7339 | 8.7221 | 8.7103 | 8.6985 | 8.6867 | 8.6749 | 8.6631 | 8.6513 | 8.6395 | 8.6277 | 8.6159 | 8.6041 | 8.5923 | 8.5805 | 8.5687 | 8.5569 | 8.5451 | 8.5333 | 8.5215 | 8.5097 | 8.4979 | 8.4861 | 8.4743 | 8.4625 | 8.4507 | 8.4389 | 8.4271 | 8.4153 | 8.4035 | 8.3917 | 8.3799 | 8.3681 | 8.3563 | 8.3445 | 8.3327 | 8.3209 | 8.3091 | 8.2973 | 8.2854 | 8.2736 | 8.2618 | 8.25 | 8.2382 | 8.2264 | 8.2146 | 8.2028 | 8.191 | 8.1792 | 8.1674 | 8.1556 | 8.1438
20 | 0.6 | 20.6 'C | 8.95 | 8.8911 | 8.8322 | 8.7734 | 8.7145 | 8.6556 | 8.5967 | 8.5378 | 8.4789 | 8.4201 | 8.3612 | 8.3023 | 8.2434 | 8.1845 | 8.1257 | 0.011776 | -1.7764e-15 | 20.6 | 8.95 | 8.9382 | 8.9264 | 8.9147 | 8.9029 | 8.8911 | 8.8793 | 8.8676 | 8.8558 | 8.844 | 8.8322 | 8.8205 | 8.8087 | 8.7969 | 8.7851 | 8.7734 | 8.7616 | 8.7498 | 8.738 | 8.7263 | 8.7145 | 8.7027 | 8.6909 | 8.6791 | 8.6674 | 8.6556 | 8.6438 | 8.632 | 8.6203 | 8.6085 | 8.5967 | 8.5849 | 8.5732 | 8.5614 | 8.5496 | 8.5378 | 8.5261 | 8.5143 | 8.5025 | 8.4907 | 8.4789 | 8.4672 | 8.4554 | 8.4436 | 8.4318 | 8.4201 | 8.4083 | 8.3965 | 8.3847 | 8.373 | 8.3612 | 8.3494 | 8.3376 | 8.3259 | 8.3141 | 8.3023 | 8.2905 | 8.2788 | 8.267 | 8.2552 | 8.2434 | 8.2316 | 8.2199 | 8.2081 | 8.1963 | 8.1845 | 8.1728 | 8.161 | 8.1492 | 8.1374 | 8.1257
20 | 0.7 | 20.7 'C | 8.93 | 8.8712 | 8.8125 | 8.7538 | 8.695 | 8.6363 | 8.5775 | 8.5188 | 8.46 | 8.4012 | 8.3425 | 8.2837 | 8.225 | 8.1662 | 8.1075 | 0.01175 | -1.7764e-15 | 20.7 | 8.93 | 8.9183 | 8.9065 | 8.8948 | 8.883 | 8.8712 | 8.8595 | 8.8477 | 8.836 | 8.8242 | 8.8125 | 8.8008 | 8.789 | 8.7773 | 8.7655 | 8.7538 | 8.742 | 8.7302 | 8.7185 | 8.7067 | 8.695 | 8.6832 | 8.6715 | 8.6598 | 8.648 | 8.6363 | 8.6245 | 8.6128 | 8.601 | 8.5892 | 8.5775 | 8.5657 | 8.554 | 8.5422 | 8.5305 | 8.5188 | 8.507 | 8.4953 | 8.4835 | 8.4718 | 8.46 | 8.4482 | 8.4365 | 8.4247 | 8.413 | 8.4012 | 8.3895 | 8.3778 | 8.366 | 8.3543 | 8.3425 | 8.3308 | 8.319 | 8.3072 | 8.2955 | 8.2837 | 8.272 | 8.2602 | 8.2485 | 8.2368 | 8.225 | 8.2133 | 8.2015 | 8.1898 | 8.178 | 8.1662 | 8.1545 | 8.1427 | 8.131 | 8.1192 | 8.1075
20 | 0.8 | 20.8 'C | 8.92 | 8.8613 | 8.8026 | 8.7439 | 8.6853 | 8.6266 | 8.5679 | 8.5092 | 8.4505 | 8.3918 | 8.3332 | 8.2745 | 8.2158 | 8.1571 | 8.0984 | 0.011737 | 7.1054e-15 | 20.8 | 8.92 | 8.9083 | 8.8965 | 8.8848 | 8.8731 | 8.8613 | 8.8496 | 8.8378 | 8.8261 | 8.8144 | 8.8026 | 8.7909 | 8.7792 | 8.7674 | 8.7557 | 8.7439 | 8.7322 | 8.7205 | 8.7087 | 8.697 | 8.6853 | 8.6735 | 8.6618 | 8.6501 | 8.6383 | 8.6266 | 8.6148 | 8.6031 | 8.5914 | 8.5796 | 8.5679 | 8.5562 | 8.5444 | 8.5327 | 8.5209 | 8.5092 | 8.4975 | 8.4857 | 8.474 | 8.4623 | 8.4505 | 8.4388 | 8.4271 | 8.4153 | 8.4036 | 8.3918 | 8.3801 | 8.3684 | 8.3566 | 8.3449 | 8.3332 | 8.3214 | 8.3097 | 8.2979 | 8.2862 | 8.2745 | 8.2627 | 8.251 | 8.2393 | 8.2275 | 8.2158 | 8.2041 | 8.1923 | 8.1806 | 8.1688 | 8.1571 | 8.1454 | 8.1336 | 8.1219 | 8.1102 | 8.0984
20 | 0.9 | 20.9 'C | 8.9 | 8.8414 | 8.7829 | 8.7243 | 8.6658 | 8.6072 | 8.5487 | 8.4901 | 8.4316 | 8.373 | 8.3145 | 8.2559 | 8.1974 | 8.1388 | 8.0803 | 0.011711 | 0 | 20.9 | 8.9 | 8.8883 | 8.8766 | 8.8649 | 8.8532 | 8.8414 | 8.8297 | 8.818 | 8.8063 | 8.7946 | 8.7829 | 8.7712 | 8.7595 | 8.7478 | 8.7361 | 8.7243 | 8.7126 | 8.7009 | 8.6892 | 8.6775 | 8.6658 | 8.6541 | 8.6424 | 8.6307 | 8.6189 | 8.6072 | 8.5955 | 8.5838 | 8.5721 | 8.5604 | 8.5487 | 8.537 | 8.5253 | 8.5136 | 8.5018 | 8.4901 | 8.4784 | 8.4667 | 8.455 | 8.4433 | 8.4316 | 8.4199 | 8.4082 | 8.3964 | 8.3847 | 8.373 | 8.3613 | 8.3496 | 8.3379 | 8.3262 | 8.3145 | 8.3028 | 8.2911 | 8.2793 | 8.2676 | 8.2559 | 8.2442 | 8.2325 | 8.2208 | 8.2091 | 8.1974 | 8.1857 | 8.1739 | 8.1622 | 8.1505 | 8.1388 | 8.1271 | 8.1154 | 8.1037 | 8.092 | 8.0803
21 | 0 | 21 'C | 8.88 | 8.8216 | 8.7632 | 8.7047 | 8.6463 | 8.5879 | 8.5295 | 8.4711 | 8.4126 | 8.3542 | 8.2958 | 8.2374 | 8.1789 | 8.1205 | 8.0621 | 0.011684 | -1.7764e-15 | 21 | 8.88 | 8.8683 | 8.8566 | 8.8449 | 8.8333 | 8.8216 | 8.8099 | 8.7982 | 8.7865 | 8.7748 | 8.7632 | 8.7515 | 8.7398 | 8.7281 | 8.7164 | 8.7047 | 8.6931 | 8.6814 | 8.6697 | 8.658 | 8.6463 | 8.6346 | 8.6229 | 8.6113 | 8.5996 | 8.5879 | 8.5762 | 8.5645 | 8.5528 | 8.5412 | 8.5295 | 8.5178 | 8.5061 | 8.4944 | 8.4827 | 8.4711 | 8.4594 | 8.4477 | 8.436 | 8.4243 | 8.4126 | 8.4009 | 8.3893 | 8.3776 | 8.3659 | 8.3542 | 8.3425 | 8.3308 | 8.3192 | 8.3075 | 8.2958 | 8.2841 | 8.2724 | 8.2607 | 8.2491 | 8.2374 | 8.2257 | 8.214 | 8.2023 | 8.1906 | 8.1789 | 8.1673 | 8.1556 | 8.1439 | 8.1322 | 8.1205 | 8.1088 | 8.0972 | 8.0855 | 8.0738 | 8.0621
21 | 0.1 | 21.1 'C | 8.86 | 8.8017 | 8.7434 | 8.6851 | 8.6268 | 8.5686 | 8.5103 | 8.452 | 8.3937 | 8.3354 | 8.2771 | 8.2188 | 8.1605 | 8.1022 | 8.0439 | 0.011658 | 3.5527e-15 | 21.1 | 8.86 | 8.8483 | 8.8367 | 8.825 | 8.8134 | 8.8017 | 8.7901 | 8.7784 | 8.7667 | 8.7551 | 8.7434 | 8.7318 | 8.7201 | 8.7084 | 8.6968 | 8.6851 | 8.6735 | 8.6618 | 8.6502 | 8.6385 | 8.6268 | 8.6152 | 8.6035 | 8.5919 | 8.5802 | 8.5686 | 8.5569 | 8.5452 | 8.5336 | 8.5219 | 8.5103 | 8.4986 | 8.4869 | 8.4753 | 8.4636 | 8.452 | 8.4403 | 8.4287 | 8.417 | 8.4053 | 8.3937 | 8.382 | 8.3704 | 8.3587 | 8.3471 | 8.3354 | 8.3237 | 8.3121 | 8.3004 | 8.2888 | 8.2771 | 8.2654 | 8.2538 | 8.2421 | 8.2305 | 8.2188 | 8.2072 | 8.1955 | 8.1838 | 8.1722 | 8.1605 | 8.1489 | 8.1372 | 8.1256 | 8.1139 | 8.1022 | 8.0906 | 8.0789 | 8.0673 | 8.0556 | 8.0439
21 | 0.2 | 21.2 'C | 8.85 | 8.7918 | 8.7336 | 8.6753 | 8.6171 | 8.5589 | 8.5007 | 8.4424 | 8.3842 | 8.326 | 8.2678 | 8.2095 | 8.1513 | 8.0931 | 8.0349 | 0.011645 | 1.7764e-15 | 21.2 | 8.85 | 8.8384 | 8.8267 | 8.8151 | 8.8034 | 8.7918 | 8.7801 | 8.7685 | 8.7568 | 8.7452 | 8.7336 | 8.7219 | 8.7103 | 8.6986 | 8.687 | 8.6753 | 8.6637 | 8.652 | 8.6404 | 8.6288 | 8.6171 | 8.6055 | 8.5938 | 8.5822 | 8.5705 | 8.5589 | 8.5472 | 8.5356 | 8.5239 | 8.5123 | 8.5007 | 8.489 | 8.4774 | 8.4657 | 8.4541 | 8.4424 | 8.4308 | 8.4191 | 8.4075 | 8.3959 | 8.3842 | 8.3726 | 8.3609 | 8.3493 | 8.3376 | 8.326 | 8.3143 | 8.3027 | 8.2911 | 8.2794 | 8.2678 | 8.2561 | 8.2445 | 8.2328 | 8.2212 | 8.2095 | 8.1979 | 8.1862 | 8.1746 | 8.163 | 8.1513 | 8.1397 | 8.128 | 8.1164 | 8.1047 | 8.0931 | 8.0814 | 8.0698 | 8.0582 | 8.0465 | 8.0349
21 | 0.3 | 21.3 'C | 8.83 | 8.7719 | 8.7138 | 8.6557 | 8.5976 | 8.5395 | 8.4814 | 8.4234 | 8.3653 | 8.3072 | 8.2491 | 8.191 | 8.1329 | 8.0748 | 8.0167 | 0.011618 | 1.2434e-14 | 21.3 | 8.83 | 8.8184 | 8.8068 | 8.7951 | 8.7835 | 8.7719 | 8.7603 | 8.7487 | 8.7371 | 8.7254 | 8.7138 | 8.7022 | 8.6906 | 8.679 | 8.6673 | 8.6557 | 8.6441 | 8.6325 | 8.6209 | 8.6092 | 8.5976 | 8.586 | 8.5744 | 8.5628 | 8.5512 | 8.5395 | 8.5279 | 8.5163 | 8.5047 | 8.4931 | 8.4814 | 8.4698 | 8.4582 | 8.4466 | 8.435 | 8.4234 | 8.4117 | 8.4001 | 8.3885 | 8.3769 | 8.3653 | 8.3536 | 8.342 | 8.3304 | 8.3188 | 8.3072 | 8.2956 | 8.2839 | 8.2723 | 8.2607 | 8.2491 | 8.2375 | 8.2258 | 8.2142 | 8.2026 | 8.191 | 8.1794 | 8.1677 | 8.1561 | 8.1445 | 8.1329 | 8.1213 | 8.1097 | 8.098 | 8.0864 | 8.0748 | 8.0632 | 8.0516 | 8.0399 | 8.0283 | 8.0167
21 | 0.4 | 21.4 'C | 8.81 | 8.752 | 8.6941 | 8.6361 | 8.5782 | 8.5202 | 8.4622 | 8.4043 | 8.3463 | 8.2884 | 8.2304 | 8.1724 | 8.1145 | 8.0565 | 7.9986 | 0.011592 | 3.5527e-15 | 21.4 | 8.81 | 8.7984 | 8.7868 | 8.7752 | 8.7636 | 8.752 | 8.7404 | 8.7289 | 8.7173 | 8.7057 | 8.6941 | 8.6825 | 8.6709 | 8.6593 | 8.6477 | 8.6361 | 8.6245 | 8.6129 | 8.6013 | 8.5898 | 8.5782 | 8.5666 | 8.555 | 8.5434 | 8.5318 | 8.5202 | 8.5086 | 8.497 | 8.4854 | 8.4738 | 8.4622 | 8.4506 | 8.4391 | 8.4275 | 8.4159 | 8.4043 | 8.3927 | 8.3811 | 8.3695 | 8.3579 | 8.3463 | 8.3347 | 8.3231 | 8.3115 | 8.2999 | 8.2884 | 8.2768 | 8.2652 | 8.2536 | 8.242 | 8.2304 | 8.2188 | 8.2072 | 8.1956 | 8.184 | 8.1724 | 8.1608 | 8.1493 | 8.1377 | 8.1261 | 8.1145 | 8.1029 | 8.0913 | 8.0797 | 8.0681 | 8.0565 | 8.0449 | 8.0333 | 8.0217 | 8.0101 | 7.9986
21 | 0.5 | 21.5 'C | 8.8 | 8.7421 | 8.6842 | 8.6263 | 8.5684 | 8.5105 | 8.4526 | 8.3947 | 8.3368 | 8.2789 | 8.2211 | 8.1632 | 8.1053 | 8.0474 | 7.9895 | 0.011579 | 3.5527e-15 | 21.5 | 8.8 | 8.7884 | 8.7768 | 8.7653 | 8.7537 | 8.7421 | 8.7305 | 8.7189 | 8.7074 | 8.6958 | 8.6842 | 8.6726 | 8.6611 | 8.6495 | 8.6379 | 8.6263 | 8.6147 | 8.6032 | 8.5916 | 8.58 | 8.5684 | 8.5568 | 8.5453 | 8.5337 | 8.5221 | 8.5105 | 8.4989 | 8.4874 | 8.4758 | 8.4642 | 8.4526 | 8.4411 | 8.4295 | 8.4179 | 8.4063 | 8.3947 | 8.3832 | 8.3716 | 8.36 | 8.3484 | 8.3368 | 8.3253 | 8.3137 | 8.3021 | 8.2905 | 8.2789 | 8.2674 | 8.2558 | 8.2442 | 8.2326 | 8.2211 | 8.2095 | 8.1979 | 8.1863 | 8.1747 | 8.1632 | 8.1516 | 8.14 | 8.1284 | 8.1168 | 8.1053 | 8.0937 | 8.0821 | 8.0705 | 8.0589 | 8.0474 | 8.0358 | 8.0242 | 8.0126 | 8.0011 | 7.9895
21 | 0.6 | 21.6 'C | 8.78 | 8.7222 | 8.6645 | 8.6067 | 8.5489 | 8.4912 | 8.4334 | 8.3757 | 8.3179 | 8.2601 | 8.2024 | 8.1446 | 8.0868 | 8.0291 | 7.9713 | 0.011553 | 0 | 21.6 | 8.78 | 8.7684 | 8.7569 | 8.7453 | 8.7338 | 8.7222 | 8.7107 | 8.6991 | 8.6876 | 8.676 | 8.6645 | 8.6529 | 8.6414 | 8.6298 | 8.6183 | 8.6067 | 8.5952 | 8.5836 | 8.5721 | 8.5605 | 8.5489 | 8.5374 | 8.5258 | 8.5143 | 8.5027 | 8.4912 | 8.4796 | 8.4681 | 8.4565 | 8.445 | 8.4334 | 8.4219 | 8.4103 | 8.3988 | 8.3872 | 8.3757 | 8.3641 | 8.3526 | 8.341 | 8.3294 | 8.3179 | 8.3063 | 8.2948 | 8.2832 | 8.2717 | 8.2601 | 8.2486 | 8.237 | 8.2255 | 8.2139 | 8.2024 | 8.1908 | 8.1793 | 8.1677 | 8.1562 | 8.1446 | 8.1331 | 8.1215 | 8.1099 | 8.0984 | 8.0868 | 8.0753 | 8.0637 | 8.0522 | 8.0406 | 8.0291 | 8.0175 | 8.006 | 7.9944 | 7.9829 | 7.9713
21 | 0.7 | 21.7 'C | 8.76 | 8.7024 | 8.6447 | 8.5871 | 8.5295 | 8.4718 | 8.4142 | 8.3566 | 8.2989 | 8.2413 | 8.1837 | 8.1261 | 8.0684 | 8.0108 | 7.9532 | 0.011526 | 3.5527e-15 | 21.7 | 8.76 | 8.7485 | 8.7369 | 8.7254 | 8.7139 | 8.7024 | 8.6908 | 8.6793 | 8.6678 | 8.6563 | 8.6447 | 8.6332 | 8.6217 | 8.6102 | 8.5986 | 8.5871 | 8.5756 | 8.5641 | 8.5525 | 8.541 | 8.5295 | 8.5179 | 8.5064 | 8.4949 | 8.4834 | 8.4718 | 8.4603 | 8.4488 | 8.4373 | 8.4257 | 8.4142 | 8.4027 | 8.3912 | 8.3796 | 8.3681 | 8.3566 | 8.3451 | 8.3335 | 8.322 | 8.3105 | 8.2989 | 8.2874 | 8.2759 | 8.2644 | 8.2528 | 8.2413 | 8.2298 | 8.2183 | 8.2067 | 8.1952 | 8.1837 | 8.1722 | 8.1606 | 8.1491 | 8.1376 | 8.1261 | 8.1145 | 8.103 | 8.0915 | 8.0799 | 8.0684 | 8.0569 | 8.0454 | 8.0338 | 8.0223 | 8.0108 | 7.9993 | 7.9877 | 7.9762 | 7.9647 | 7.9532
21 | 0.8 | 21.8 'C | 8.74 | 8.6825 | 8.625 | 8.5675 | 8.51 | 8.4525 | 8.395 | 8.3375 | 8.28 | 8.2225 | 8.165 | 8.1075 | 8.05 | 7.9925 | 7.935 | 0.0115 | 1.7764e-15 | 21.8 | 8.74 | 8.7285 | 8.717 | 8.7055 | 8.694 | 8.6825 | 8.671 | 8.6595 | 8.648 | 8.6365 | 8.625 | 8.6135 | 8.602 | 8.5905 | 8.579 | 8.5675 | 8.556 | 8.5445 | 8.533 | 8.5215 | 8.51 | 8.4985 | 8.487 | 8.4755 | 8.464 | 8.4525 | 8.441 | 8.4295 | 8.418 | 8.4065 | 8.395 | 8.3835 | 8.372 | 8.3605 | 8.349 | 8.3375 | 8.326 | 8.3145 | 8.303 | 8.2915 | 8.28 | 8.2685 | 8.257 | 8.2455 | 8.234 | 8.2225 | 8.211 | 8.1995 | 8.188 | 8.1765 | 8.165 | 8.1535 | 8.142 | 8.1305 | 8.119 | 8.1075 | 8.096 | 8.0845 | 8.073 | 8.0615 | 8.05 | 8.0385 | 8.027 | 8.0155 | 8.004 | 7.9925 | 7.981 | 7.9695 | 7.958 | 7.9465 | 7.935
21 | 0.9 | 21.9 'C | 8.73 | 8.6726 | 8.6151 | 8.5577 | 8.5003 | 8.4428 | 8.3854 | 8.328 | 8.2705 | 8.2131 | 8.1557 | 8.0982 | 8.0408 | 7.9834 | 7.9259 | 0.011487 | 0 | 21.9 | 8.73 | 8.7185 | 8.707 | 8.6955 | 8.6841 | 8.6726 | 8.6611 | 8.6496 | 8.6381 | 8.6266 | 8.6151 | 8.6036 | 8.5922 | 8.5807 | 8.5692 | 8.5577 | 8.5462 | 8.5347 | 8.5232 | 8.5117 | 8.5003 | 8.4888 | 8.4773 | 8.4658 | 8.4543 | 8.4428 | 8.4313 | 8.4199 | 8.4084 | 8.3969 | 8.3854 | 8.3739 | 8.3624 | 8.3509 | 8.3394 | 8.328 | 8.3165 | 8.305 | 8.2935 | 8.282 | 8.2705 | 8.259 | 8.2476 | 8.2361 | 8.2246 | 8.2131 | 8.2016 | 8.1901 | 8.1786 | 8.1671 | 8.1557 | 8.1442 | 8.1327 | 8.1212 | 8.1097 | 8.0982 | 8.0867 | 8.0753 | 8.0638 | 8.0523 | 8.0408 | 8.0293 | 8.0178 | 8.0063 | 7.9948 | 7.9834 | 7.9719 | 7.9604 | 7.9489 | 7.9374 | 7.9259
22 | 0 | 22 'C | 8.71 | 8.6527 | 8.5954 | 8.5381 | 8.4808 | 8.4235 | 8.3662 | 8.3089 | 8.2516 | 8.1943 | 8.137 | 8.0797 | 8.0224 | 7.9651 | 7.9078 | 0.011461 | 3.5527e-15 | 22 | 8.71 | 8.6985 | 8.6871 | 8.6756 | 8.6642 | 8.6527 | 8.6412 | 8.6298 | 8.6183 | 8.6069 | 8.5954 | 8.5839 | 8.5725 | 8.561 | 8.5496 | 8.5381 | 8.5266 | 8.5152 | 8.5037 | 8.4923 | 8.4808 | 8.4693 | 8.4579 | 8.4464 | 8.4349 | 8.4235 | 8.412 | 8.4006 | 8.3891 | 8.3776 | 8.3662 | 8.3547 | 8.3433 | 8.3318 | 8.3203 | 8.3089 | 8.2974 | 8.286 | 8.2745 | 8.263 | 8.2516 | 8.2401 | 8.2287 | 8.2172 | 8.2057 | 8.1943 | 8.1828 | 8.1714 | 8.1599 | 8.1484 | 8.137 | 8.1255 | 8.1141 | 8.1026 | 8.0911 | 8.0797 | 8.0682 | 8.0567 | 8.0453 | 8.0338 | 8.0224 | 8.0109 | 7.9994 | 7.988 | 7.9765 | 7.9651 | 7.9536 | 7.9421 | 7.9307 | 7.9192 | 7.9078
22 | 0.1 | 22.1 'C | 8.69 | 8.6328 | 8.5757 | 8.5185 | 8.4613 | 8.4041 | 8.347 | 8.2898 | 8.2326 | 8.1755 | 8.1183 | 8.0611 | 8.0039 | 7.9468 | 7.8896 | 0.011434 | 8.8818e-15 | 22.1 | 8.69 | 8.6786 | 8.6671 | 8.6557 | 8.6443 | 8.6328 | 8.6214 | 8.61 | 8.5985 | 8.5871 | 8.5757 | 8.5642 | 8.5528 | 8.5414 | 8.5299 | 8.5185 | 8.5071 | 8.4956 | 8.4842 | 8.4727 | 8.4613 | 8.4499 | 8.4384 | 8.427 | 8.4156 | 8.4041 | 8.3927 | 8.3813 | 8.3698 | 8.3584 | 8.347 | 8.3355 | 8.3241 | 8.3127 | 8.3012 | 8.2898 | 8.2784 | 8.2669 | 8.2555 | 8.2441 | 8.2326 | 8.2212 | 8.2098 | 8.1983 | 8.1869 | 8.1755 | 8.164 | 8.1526 | 8.1412 | 8.1297 | 8.1183 | 8.1069 | 8.0954 | 8.084 | 8.0726 | 8.0611 | 8.0497 | 8.0382 | 8.0268 | 8.0154 | 8.0039 | 7.9925 | 7.9811 | 7.9696 | 7.9582 | 7.9468 | 7.9353 | 7.9239 | 7.9125 | 7.901 | 7.8896
22 | 0.2 | 22.2 'C | 8.68 | 8.6229 | 8.5658 | 8.5087 | 8.4516 | 8.3945 | 8.3374 | 8.2803 | 8.2232 | 8.1661 | 8.1089 | 8.0518 | 7.9947 | 7.9376 | 7.8805 | 0.011421 | -1.7764e-15 | 22.2 | 8.68 | 8.6686 | 8.6572 | 8.6457 | 8.6343 | 8.6229 | 8.6115 | 8.6001 | 8.5886 | 8.5772 | 8.5658 | 8.5544 | 8.5429 | 8.5315 | 8.5201 | 8.5087 | 8.4973 | 8.4858 | 8.4744 | 8.463 | 8.4516 | 8.4402 | 8.4287 | 8.4173 | 8.4059 | 8.3945 | 8.3831 | 8.3716 | 8.3602 | 8.3488 | 8.3374 | 8.3259 | 8.3145 | 8.3031 | 8.2917 | 8.2803 | 8.2688 | 8.2574 | 8.246 | 8.2346 | 8.2232 | 8.2117 | 8.2003 | 8.1889 | 8.1775 | 8.1661 | 8.1546 | 8.1432 | 8.1318 | 8.1204 | 8.1089 | 8.0975 | 8.0861 | 8.0747 | 8.0633 | 8.0518 | 8.0404 | 8.029 | 8.0176 | 8.0062 | 7.9947 | 7.9833 | 7.9719 | 7.9605 | 7.9491 | 7.9376 | 7.9262 | 7.9148 | 7.9034 | 7.8919 | 7.8805
22 | 0.3 | 22.3 'C | 8.66 | 8.603 | 8.5461 | 8.4891 | 8.4321 | 8.3751 | 8.3182 | 8.2612 | 8.2042 | 8.1472 | 8.0903 | 8.0333 | 7.9763 | 7.9193 | 7.8624 | 0.011395 | 5.3291e-15 | 22.3 | 8.66 | 8.6486 | 8.6372 | 8.6258 | 8.6144 | 8.603 | 8.5916 | 8.5802 | 8.5688 | 8.5574 | 8.5461 | 8.5347 | 8.5233 | 8.5119 | 8.5005 | 8.4891 | 8.4777 | 8.4663 | 8.4549 | 8.4435 | 8.4321 | 8.4207 | 8.4093 | 8.3979 | 8.3865 | 8.3751 | 8.3637 | 8.3523 | 8.3409 | 8.3296 | 8.3182 | 8.3068 | 8.2954 | 8.284 | 8.2726 | 8.2612 | 8.2498 | 8.2384 | 8.227 | 8.2156 | 8.2042 | 8.1928 | 8.1814 | 8.17 | 8.1586 | 8.1472 | 8.1358 | 8.1244 | 8.1131 | 8.1017 | 8.0903 | 8.0789 | 8.0675 | 8.0561 | 8.0447 | 8.0333 | 8.0219 | 8.0105 | 7.9991 | 7.9877 | 7.9763 | 7.9649 | 7.9535 | 7.9421 | 7.9307 | 7.9193 | 7.9079 | 7.8966 | 7.8852 | 7.8738 | 7.8624
22 | 0.4 | 22.4 'C | 8.65 | 8.5931 | 8.5362 | 8.4793 | 8.4224 | 8.3655 | 8.3086 | 8.2516 | 8.1947 | 8.1378 | 8.0809 | 8.024 | 7.9671 | 7.9102 | 7.8533 | 0.011382 | 3.5527e-15 | 22.4 | 8.65 | 8.6386 | 8.6272 | 8.6159 | 8.6045 | 8.5931 | 8.5817 | 8.5703 | 8.5589 | 8.5476 | 8.5362 | 8.5248 | 8.5134 | 8.502 | 8.4907 | 8.4793 | 8.4679 | 8.4565 | 8.4451 | 8.4337 | 8.4224 | 8.411 | 8.3996 | 8.3882 | 8.3768 | 8.3655 | 8.3541 | 8.3427 | 8.3313 | 8.3199 | 8.3086 | 8.2972 | 8.2858 | 8.2744 | 8.263 | 8.2516 | 8.2403 | 8.2289 | 8.2175 | 8.2061 | 8.1947 | 8.1834 | 8.172 | 8.1606 | 8.1492 | 8.1378 | 8.1264 | 8.1151 | 8.1037 | 8.0923 | 8.0809 | 8.0695 | 8.0582 | 8.0468 | 8.0354 | 8.024 | 8.0126 | 8.0013 | 7.9899 | 7.9785 | 7.9671 | 7.9557 | 7.9443 | 7.933 | 7.9216 | 7.9102 | 7.8988 | 7.8874 | 7.8761 | 7.8647 | 7.8533
22 | 0.5 | 22.5 'C | 8.63 | 8.5732 | 8.5164 | 8.4597 | 8.4029 | 8.3461 | 8.2893 | 8.2326 | 8.1758 | 8.119 | 8.0622 | 8.0055 | 7.9487 | 7.8919 | 7.8351 | 0.011355 | 7.1054e-15 | 22.5 | 8.63 | 8.6186 | 8.6073 | 8.5959 | 8.5846 | 8.5732 | 8.5619 | 8.5505 | 8.5392 | 8.5278 | 8.5164 | 8.5051 | 8.4937 | 8.4824 | 8.471 | 8.4597 | 8.4483 | 8.437 | 8.4256 | 8.4142 | 8.4029 | 8.3915 | 8.3802 | 8.3688 | 8.3575 | 8.3461 | 8.3348 | 8.3234 | 8.3121 | 8.3007 | 8.2893 | 8.278 | 8.2666 | 8.2553 | 8.2439 | 8.2326 | 8.2212 | 8.2099 | 8.1985 | 8.1871 | 8.1758 | 8.1644 | 8.1531 | 8.1417 | 8.1304 | 8.119 | 8.1077 | 8.0963 | 8.0849 | 8.0736 | 8.0622 | 8.0509 | 8.0395 | 8.0282 | 8.0168 | 8.0055 | 7.9941 | 7.9828 | 7.9714 | 7.96 | 7.9487 | 7.9373 | 7.926 | 7.9146 | 7.9033 | 7.8919 | 7.8806 | 7.8692 | 7.8578 | 7.8465 | 7.8351
22 | 0.6 | 22.6 'C | 8.61 | 8.5534 | 8.4967 | 8.4401 | 8.3834 | 8.3268 | 8.2701 | 8.2135 | 8.1568 | 8.1002 | 8.0436 | 7.9869 | 7.9303 | 7.8736 | 7.817 | 0.011329 | 0 | 22.6 | 8.61 | 8.5987 | 8.5873 | 8.576 | 8.5647 | 8.5534 | 8.542 | 8.5307 | 8.5194 | 8.508 | 8.4967 | 8.4854 | 8.4741 | 8.4627 | 8.4514 | 8.4401 | 8.4287 | 8.4174 | 8.4061 | 8.3948 | 8.3834 | 8.3721 | 8.3608 | 8.3494 | 8.3381 | 8.3268 | 8.3154 | 8.3041 | 8.2928 | 8.2815 | 8.2701 | 8.2588 | 8.2475 | 8.2361 | 8.2248 | 8.2135 | 8.2022 | 8.1908 | 8.1795 | 8.1682 | 8.1568 | 8.1455 | 8.1342 | 8.1229 | 8.1115 | 8.1002 | 8.0889 | 8.0775 | 8.0662 | 8.0549 | 8.0436 | 8.0322 | 8.0209 | 8.0096 | 7.9982 | 7.9869 | 7.9756 | 7.9642 | 7.9529 | 7.9416 | 7.9303 | 7.9189 | 7.9076 | 7.8963 | 7.8849 | 7.8736 | 7.8623 | 7.851 | 7.8396 | 7.8283 | 7.817
22 | 0.7 | 22.7 'C | 8.6 | 8.5434 | 8.4868 | 8.4303 | 8.3737 | 8.3171 | 8.2605 | 8.2039 | 8.1474 | 8.0908 | 8.0342 | 7.9776 | 7.9211 | 7.8645 | 7.8079 | 0.011316 | 0 | 22.7 | 8.6 | 8.5887 | 8.5774 | 8.5661 | 8.5547 | 8.5434 | 8.5321 | 8.5208 | 8.5095 | 8.4982 | 8.4868 | 8.4755 | 8.4642 | 8.4529 | 8.4416 | 8.4303 | 8.4189 | 8.4076 | 8.3963 | 8.385 | 8.3737 | 8.3624 | 8.3511 | 8.3397 | 8.3284 | 8.3171 | 8.3058 | 8.2945 | 8.2832 | 8.2718 | 8.2605 | 8.2492 | 8.2379 | 8.2266 | 8.2153 | 8.2039 | 8.1926 | 8.1813 | 8.17 | 8.1587 | 8.1474 | 8.1361 | 8.1247 | 8.1134 | 8.1021 | 8.0908 | 8.0795 | 8.0682 | 8.0568 | 8.0455 | 8.0342 | 8.0229 | 8.0116 | 8.0003 | 7.9889 | 7.9776 | 7.9663 | 7.955 | 7.9437 | 7.9324 | 7.9211 | 7.9097 | 7.8984 | 7.8871 | 7.8758 | 7.8645 | 7.8532 | 7.8418 | 7.8305 | 7.8192 | 7.8079
22 | 0.8 | 22.8 'C | 8.58 | 8.5236 | 8.4671 | 8.4107 | 8.3542 | 8.2978 | 8.2413 | 8.1849 | 8.1284 | 8.072 | 8.0155 | 7.9591 | 7.9026 | 7.8462 | 7.7897 | 0.011289 | 5.3291e-15 | 22.8 | 8.58 | 8.5687 | 8.5574 | 8.5461 | 8.5348 | 8.5236 | 8.5123 | 8.501 | 8.4897 | 8.4784 | 8.4671 | 8.4558 | 8.4445 | 8.4332 | 8.4219 | 8.4107 | 8.3994 | 8.3881 | 8.3768 | 8.3655 | 8.3542 | 8.3429 | 8.3316 | 8.3203 | 8.3091 | 8.2978 | 8.2865 | 8.2752 | 8.2639 | 8.2526 | 8.2413 | 8.23 | 8.2187 | 8.2074 | 8.1962 | 8.1849 | 8.1736 | 8.1623 | 8.151 | 8.1397 | 8.1284 | 8.1171 | 8.1058 | 8.0946 | 8.0833 | 8.072 | 8.0607 | 8.0494 | 8.0381 | 8.0268 | 8.0155 | 8.0042 | 7.9929 | 7.9817 | 7.9704 | 7.9591 | 7.9478 | 7.9365 | 7.9252 | 7.9139 | 7.9026 | 7.8913 | 7.8801 | 7.8688 | 7.8575 | 7.8462 | 7.8349 | 7.8236 | 7.8123 | 7.801 | 7.7897
22 | 0.9 | 22.9 'C | 8.57 | 8.5136 | 8.4572 | 8.4009 | 8.3445 | 8.2881 | 8.2317 | 8.1753 | 8.1189 | 8.0626 | 8.0062 | 7.9498 | 7.8934 | 7.837 | 7.7807 | 0.011276 | 0 | 22.9 | 8.57 | 8.5587 | 8.5474 | 8.5362 | 8.5249 | 8.5136 | 8.5023 | 8.4911 | 8.4798 | 8.4685 | 8.4572 | 8.446 | 8.4347 | 8.4234 | 8.4121 | 8.4009 | 8.3896 | 8.3783 | 8.367 | 8.3558 | 8.3445 | 8.3332 | 8.3219 | 8.3106 | 8.2994 | 8.2881 | 8.2768 | 8.2655 | 8.2543 | 8.243 | 8.2317 | 8.2204 | 8.2092 | 8.1979 | 8.1866 | 8.1753 | 8.1641 | 8.1528 | 8.1415 | 8.1302 | 8.1189 | 8.1077 | 8.0964 | 8.0851 | 8.0738 | 8.0626 | 8.0513 | 8.04 | 8.0287 | 8.0175 | 8.0062 | 7.9949 | 7.9836 | 7.9724 | 7.9611 | 7.9498 | 7.9385 | 7.9272 | 7.916 | 7.9047 | 7.8934 | 7.8821 | 7.8709 | 7.8596 | 7.8483 | 7.837 | 7.8258 | 7.8145 | 7.8032 | 7.7919 | 7.7807
23 | 0 | 23 'C | 8.55 | 8.4938 | 8.4375 | 8.3812 | 8.325 | 8.2688 | 8.2125 | 8.1562 | 8.1 | 8.0437 | 7.9875 | 7.9313 | 7.875 | 7.8187 | 7.7625 | 0.01125 | 5.3291e-15 | 23 | 8.55 | 8.5388 | 8.5275 | 8.5162 | 8.505 | 8.4938 | 8.4825 | 8.4712 | 8.46 | 8.4488 | 8.4375 | 8.4262 | 8.415 | 8.4038 | 8.3925 | 8.3812 | 8.37 | 8.3588 | 8.3475 | 8.3362 | 8.325 | 8.3138 | 8.3025 | 8.2912 | 8.28 | 8.2688 | 8.2575 | 8.2462 | 8.235 | 8.2238 | 8.2125 | 8.2012 | 8.19 | 8.1788 | 8.1675 | 8.1562 | 8.145 | 8.1337 | 8.1225 | 8.1113 | 8.1 | 8.0887 | 8.0775 | 8.0663 | 8.055 | 8.0437 | 8.0325 | 8.0213 | 8.01 | 7.9988 | 7.9875 | 7.9763 | 7.965 | 7.9538 | 7.9425 | 7.9313 | 7.92 | 7.9088 | 7.8975 | 7.8863 | 7.875 | 7.8637 | 7.8525 | 7.8412 | 7.83 | 7.8187 | 7.8075 | 7.7962 | 7.785 | 7.7737 | 7.7625
23 | 0.1 | 23.1 'C | 8.53 | 8.4739 | 8.4178 | 8.3616 | 8.3055 | 8.2494 | 8.1933 | 8.1372 | 8.0811 | 8.0249 | 7.9688 | 7.9127 | 7.8566 | 7.8005 | 7.7443 | 0.011224 | 1.7764e-15 | 23.1 | 8.53 | 8.5188 | 8.5076 | 8.4963 | 8.4851 | 8.4739 | 8.4627 | 8.4514 | 8.4402 | 8.429 | 8.4178 | 8.4065 | 8.3953 | 8.3841 | 8.3729 | 8.3616 | 8.3504 | 8.3392 | 8.328 | 8.3168 | 8.3055 | 8.2943 | 8.2831 | 8.2719 | 8.2606 | 8.2494 | 8.2382 | 8.227 | 8.2157 | 8.2045 | 8.1933 | 8.1821 | 8.1708 | 8.1596 | 8.1484 | 8.1372 | 8.1259 | 8.1147 | 8.1035 | 8.0923 | 8.0811 | 8.0698 | 8.0586 | 8.0474 | 8.0362 | 8.0249 | 8.0137 | 8.0025 | 7.9913 | 7.98 | 7.9688 | 7.9576 | 7.9464 | 7.9351 | 7.9239 | 7.9127 | 7.9015 | 7.8902 | 7.879 | 7.8678 | 7.8566 | 7.8454 | 7.8341 | 7.8229 | 7.8117 | 7.8005 | 7.7892 | 7.778 | 7.7668 | 7.7556 | 7.7443
23 | 0.2 | 23.2 'C | 8.52 | 8.4639 | 8.4079 | 8.3518 | 8.2958 | 8.2397 | 8.1837 | 8.1276 | 8.0716 | 8.0155 | 7.9595 | 7.9034 | 7.8474 | 7.7913 | 7.7353 | 0.011211 | 0 | 23.2 | 8.52 | 8.5088 | 8.4976 | 8.4864 | 8.4752 | 8.4639 | 8.4527 | 8.4415 | 8.4303 | 8.4191 | 8.4079 | 8.3967 | 8.3855 | 8.3743 | 8.3631 | 8.3518 | 8.3406 | 8.3294 | 8.3182 | 8.307 | 8.2958 | 8.2846 | 8.2734 | 8.2622 | 8.2509 | 8.2397 | 8.2285 | 8.2173 | 8.2061 | 8.1949 | 8.1837 | 8.1725 | 8.1613 | 8.1501 | 8.1388 | 8.1276 | 8.1164 | 8.1052 | 8.094 | 8.0828 | 8.0716 | 8.0604 | 8.0492 | 8.0379 | 8.0267 | 8.0155 | 8.0043 | 7.9931 | 7.9819 | 7.9707 | 7.9595 | 7.9483 | 7.9371 | 7.9258 | 7.9146 | 7.9034 | 7.8922 | 7.881 | 7.8698 | 7.8586 | 7.8474 | 7.8362 | 7.8249 | 7.8137 | 7.8025 | 7.7913 | 7.7801 | 7.7689 | 7.7577 | 7.7465 | 7.7353
23 | 0.3 | 23.3 'C | 8.5 | 8.4441 | 8.3882 | 8.3322 | 8.2763 | 8.2204 | 8.1645 | 8.1086 | 8.0526 | 7.9967 | 7.9408 | 7.8849 | 7.8289 | 7.773 | 7.7171 | 0.011184 | -1.7764e-15 | 23.3 | 8.5 | 8.4888 | 8.4776 | 8.4664 | 8.4553 | 8.4441 | 8.4329 | 8.4217 | 8.4105 | 8.3993 | 8.3882 | 8.377 | 8.3658 | 8.3546 | 8.3434 | 8.3322 | 8.3211 | 8.3099 | 8.2987 | 8.2875 | 8.2763 | 8.2651 | 8.2539 | 8.2428 | 8.2316 | 8.2204 | 8.2092 | 8.198 | 8.1868 | 8.1757 | 8.1645 | 8.1533 | 8.1421 | 8.1309 | 8.1197 | 8.1086 | 8.0974 | 8.0862 | 8.075 | 8.0638 | 8.0526 | 8.0414 | 8.0303 | 8.0191 | 8.0079 | 7.9967 | 7.9855 | 7.9743 | 7.9632 | 7.952 | 7.9408 | 7.9296 | 7.9184 | 7.9072 | 7.8961 | 7.8849 | 7.8737 | 7.8625 | 7.8513 | 7.8401 | 7.8289 | 7.8178 | 7.8066 | 7.7954 | 7.7842 | 7.773 | 7.7618 | 7.7507 | 7.7395 | 7.7283 | 7.7171
23 | 0.4 | 23.4 'C | 8.49 | 8.4341 | 8.3783 | 8.3224 | 8.2666 | 8.2107 | 8.1549 | 8.099 | 8.0432 | 7.9873 | 7.9314 | 7.8756 | 7.8197 | 7.7639 | 7.708 | 0.011171 | 3.5527e-15 | 23.4 | 8.49 | 8.4788 | 8.4677 | 8.4565 | 8.4453 | 8.4341 | 8.423 | 8.4118 | 8.4006 | 8.3895 | 8.3783 | 8.3671 | 8.3559 | 8.3448 | 8.3336 | 8.3224 | 8.3113 | 8.3001 | 8.2889 | 8.2777 | 8.2666 | 8.2554 | 8.2442 | 8.2331 | 8.2219 | 8.2107 | 8.1996 | 8.1884 | 8.1772 | 8.166 | 8.1549 | 8.1437 | 8.1325 | 8.1214 | 8.1102 | 8.099 | 8.0878 | 8.0767 | 8.0655 | 8.0543 | 8.0432 | 8.032 | 8.0208 | 8.0096 | 7.9985 | 7.9873 | 7.9761 | 7.965 | 7.9538 | 7.9426 | 7.9314 | 7.9203 | 7.9091 | 7.8979 | 7.8868 | 7.8756 | 7.8644 | 7.8533 | 7.8421 | 7.8309 | 7.8197 | 7.8086 | 7.7974 | 7.7862 | 7.7751 | 7.7639 | 7.7527 | 7.7415 | 7.7304 | 7.7192 | 7.708
23 | 0.5 | 23.5 'C | 8.47 | 8.4143 | 8.3586 | 8.3028 | 8.2471 | 8.1914 | 8.1357 | 8.0799 | 8.0242 | 7.9685 | 7.9128 | 7.857 | 7.8013 | 7.7456 | 7.6899 | 0.011145 | 5.3291e-15 | 23.5 | 8.47 | 8.4589 | 8.4477 | 8.4366 | 8.4254 | 8.4143 | 8.4031 | 8.392 | 8.3808 | 8.3697 | 8.3586 | 8.3474 | 8.3363 | 8.3251 | 8.314 | 8.3028 | 8.2917 | 8.2805 | 8.2694 | 8.2583 | 8.2471 | 8.236 | 8.2248 | 8.2137 | 8.2025 | 8.1914 | 8.1802 | 8.1691 | 8.1579 | 8.1468 | 8.1357 | 8.1245 | 8.1134 | 8.1022 | 8.0911 | 8.0799 | 8.0688 | 8.0576 | 8.0465 | 8.0354 | 8.0242 | 8.0131 | 8.0019 | 7.9908 | 7.9796 | 7.9685 | 7.9573 | 7.9462 | 7.9351 | 7.9239 | 7.9128 | 7.9016 | 7.8905 | 7.8793 | 7.8682 | 7.857 | 7.8459 | 7.8347 | 7.8236 | 7.8125 | 7.8013 | 7.7902 | 7.779 | 7.7679 | 7.7567 | 7.7456 | 7.7344 | 7.7233 | 7.7122 | 7.701 | 7.6899
23 | 0.6 | 23.6 'C | 8.45 | 8.3944 | 8.3388 | 8.2832 | 8.2276 | 8.172 | 8.1164 | 8.0609 | 8.0053 | 7.9497 | 7.8941 | 7.8385 | 7.7829 | 7.7273 | 7.6717 | 0.011118 | 0 | 23.6 | 8.45 | 8.4389 | 8.4278 | 8.4166 | 8.4055 | 8.3944 | 8.3833 | 8.3722 | 8.3611 | 8.3499 | 8.3388 | 8.3277 | 8.3166 | 8.3055 | 8.2943 | 8.2832 | 8.2721 | 8.261 | 8.2499 | 8.2387 | 8.2276 | 8.2165 | 8.2054 | 8.1943 | 8.1832 | 8.172 | 8.1609 | 8.1498 | 8.1387 | 8.1276 | 8.1164 | 8.1053 | 8.0942 | 8.0831 | 8.072 | 8.0609 | 8.0497 | 8.0386 | 8.0275 | 8.0164 | 8.0053 | 7.9941 | 7.983 | 7.9719 | 7.9608 | 7.9497 | 7.9386 | 7.9274 | 7.9163 | 7.9052 | 7.8941 | 7.883 | 7.8718 | 7.8607 | 7.8496 | 7.8385 | 7.8274 | 7.8163 | 7.8051 | 7.794 | 7.7829 | 7.7718 | 7.7607 | 7.7495 | 7.7384 | 7.7273 | 7.7162 | 7.7051 | 7.6939 | 7.6828 | 7.6717
23 | 0.7 | 23.7 'C | 8.44 | 8.3845 | 8.3289 | 8.2734 | 8.2179 | 8.1624 | 8.1068 | 8.0513 | 7.9958 | 7.9403 | 7.8847 | 7.8292 | 7.7737 | 7.7182 | 7.6626 | 0.011105 | 5.3291e-15 | 23.7 | 8.44 | 8.4289 | 8.4178 | 8.4067 | 8.3956 | 8.3845 | 8.3734 | 8.3623 | 8.3512 | 8.3401 | 8.3289 | 8.3178 | 8.3067 | 8.2956 | 8.2845 | 8.2734 | 8.2623 | 8.2512 | 8.2401 | 8.229 | 8.2179 | 8.2068 | 8.1957 | 8.1846 | 8.1735 | 8.1624 | 8.1513 | 8.1402 | 8.1291 | 8.1179 | 8.1068 | 8.0957 | 8.0846 | 8.0735 | 8.0624 | 8.0513 | 8.0402 | 8.0291 | 8.018 | 8.0069 | 7.9958 | 7.9847 | 7.9736 | 7.9625 | 7.9514 | 7.9403 | 7.9292 | 7.9181 | 7.9069 | 7.8958 | 7.8847 | 7.8736 | 7.8625 | 7.8514 | 7.8403 | 7.8292 | 7.8181 | 7.807 | 7.7959 | 7.7848 | 7.7737 | 7.7626 | 7.7515 | 7.7404 | 7.7293 | 7.7182 | 7.7071 | 7.6959 | 7.6848 | 7.6737 | 7.6626
23 | 0.8 | 23.8 'C | 8.42 | 8.3646 | 8.3092 | 8.2538 | 8.1984 | 8.143 | 8.0876 | 8.0322 | 7.9768 | 7.9214 | 7.8661 | 7.8107 | 7.7553 | 7.6999 | 7.6445 | 0.011079 | 1.7764e-15 | 23.8 | 8.42 | 8.4089 | 8.3978 | 8.3868 | 8.3757 | 8.3646 | 8.3535 | 8.3424 | 8.3314 | 8.3203 | 8.3092 | 8.2981 | 8.2871 | 8.276 | 8.2649 | 8.2538 | 8.2427 | 8.2317 | 8.2206 | 8.2095 | 8.1984 | 8.1873 | 8.1763 | 8.1652 | 8.1541 | 8.143 | 8.1319 | 8.1209 | 8.1098 | 8.0987 | 8.0876 | 8.0766 | 8.0655 | 8.0544 | 8.0433 | 8.0322 | 8.0212 | 8.0101 | 7.999 | 7.9879 | 7.9768 | 7.9658 | 7.9547 | 7.9436 | 7.9325 | 7.9214 | 7.9104 | 7.8993 | 7.8882 | 7.8771 | 7.8661 | 7.855 | 7.8439 | 7.8328 | 7.8217 | 7.8107 | 7.7996 | 7.7885 | 7.7774 | 7.7663 | 7.7553 | 7.7442 | 7.7331 | 7.722 | 7.7109 | 7.6999 | 7.6888 | 7.6777 | 7.6666 | 7.6556 | 7.6445
23 | 0.9 | 23.9 'C | 8.41 | 8.3547 | 8.2993 | 8.244 | 8.1887 | 8.1334 | 8.078 | 8.0227 | 7.9674 | 7.912 | 7.8567 | 7.8014 | 7.7461 | 7.6907 | 7.6354 | 0.011066 | 1.7764e-15 | 23.9 | 8.41 | 8.3989 | 8.3879 | 8.3768 | 8.3657 | 8.3547 | 8.3436 | 8.3325 | 8.3215 | 8.3104 | 8.2993 | 8.2883 | 8.2772 | 8.2661 | 8.2551 | 8.244 | 8.2329 | 8.2219 | 8.2108 | 8.1997 | 8.1887 | 8.1776 | 8.1666 | 8.1555 | 8.1444 | 8.1334 | 8.1223 | 8.1112 | 8.1002 | 8.0891 | 8.078 | 8.067 | 8.0559 | 8.0448 | 8.0338 | 8.0227 | 8.0116 | 8.0006 | 7.9895 | 7.9784 | 7.9674 | 7.9563 | 7.9452 | 7.9342 | 7.9231 | 7.912 | 7.901 | 7.8899 | 7.8788 | 7.8678 | 7.8567 | 7.8456 | 7.8346 | 7.8235 | 7.8124 | 7.8014 | 7.7903 | 7.7793 | 7.7682 | 7.7571 | 7.7461 | 7.735 | 7.7239 | 7.7129 | 7.7018 | 7.6907 | 7.6797 | 7.6686 | 7.6575 | 7.6465 | 7.6354

## Rev Record (hidden)
Revision Record
Date | Initials | Changes
2023-06-30 | ARF | Added tab for revision record. Added revision date to Pressure Verification tab. Editorial revisions.
2023-11-28 | ARF | In sampling info table, added 3 more rows, added column for pH of adjusted samples, and fixed conditional formatting for temperature and res. Cl cells. Also added adjusted pH range to QC table and added comment at the bottom to check res. Cl if samples were chlorinated. Added equations so that if sample IDs are added to sampling info table, they will automatically populate below, but all are still editable.
2023-12-04 | ARF | Corrected conditionally formatting for supersaturation to reference the correct cell.