## Annual
Annual Pipette Calibration (gravimetric)*
Date |  | Analyst |  | Balance
Temp. (C) |  | Z-Factor | #N/A | Serial #
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Corrective Action Due To Failures?
Acceptance Criteria
% Inaccuracy must be less than 2.00
No replicate may be greater than 2% from true volume.
%CV must be less than 1.00
*Created by the Wisconsin State Laboratory of Hygiene
15 | 1.002
15.5 | 1.002
16 | 1.0021
16.5 | 1.0022
17 | 1.0023
17.5 | 1.0024
18 | 1.0025
18.5 | 1.0026
19 | 1.0027
19.5 | 1.0028
20 | 1.0029
20.5 | 1.003
21 | 1.0031
21.5 | 1.0032
22 | 1.0033
22.5 | 1.0034
23 | 1.0035
23.5 | 1.0036
24 | 1.0037
24.5 | 1.0038
25 | 1.0039

## Quarterly
Quarterly Pipette Calibration (gravimetric)*
Date |  | Analyst |  | Balance
Temp. (C) |  | Z-Factor | #N/A | Serial #
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Corrective Action Due To Failures?
Acceptance Criteria
% Inaccuracy must be less than 2.00
No replicate may be greater than 2% from true volume.
%CV must be less than 1.00
*Created by the Wisconsin State Laboratory of Hygiene
15 | 1.002
15.5 | 1.002
16 | 1.0021
16.5 | 1.0022
17 | 1.0023
17.5 | 1.0024
18 | 1.0025
18.5 | 1.0026
19 | 1.0027
19.5 | 1.0028
20 | 1.0029
20.5 | 1.003
21 | 1.0031
21.5 | 1.0032
22 | 1.0033
22.5 | 1.0034
23 | 1.0035
23.5 | 1.0036
24 | 1.0037
24.5 | 1.0038
25 | 1.0039